511. Cyprodinil

HPLC method

CIPAC Collaborative Trial
according to 
CIPAC Information Sheet No 278

##### S. Völlmin

Syngenta Crop Protection Münchwilen AG
CH-4333 Münchwilen

- Switzerland

##### February 2008

### CYPRODINIL technical

511/TC/M/

- ISO common name:	Cyprodinil

Chemical name:   	 2-pyrimidinamine, 4-cyclopropyl-6-methyl-N-phenyl-

CAS-Number:  	 121552-61-2

Molecular mass: 	 225.3

Empirical formula:	 C14H15N3

1 Sampling

Take at least 50 g.

## 2 Identity test

2.1 HPLC Use the reversed phase HPLC method 3.1 described below. The relative retention time of the cyprodinil peak in the sample solution should not deviate by more than 2% from that of the calibration solution.

2.2 GC The relative retention time of the cyprodinil peak in the sample solution should not deviate more than 2% from that the calibration solution.

REAGENT

Acetone analytical grade

APPARATUS

Gas chromatography equipped with flame ionisation detector.

Wide-bore column fused silica, 15 m x 0.53 (i.d.) mm, film thickness: 1.5 m, coated with

5 % Phenyl and 95 % dimethylpolysiloxane (DB-5 or equivalent)

Data system

PROCEDURE

(a) Gas chromatographic conditions (typical):

Column			fused silica, 15 m x 0.53 (i.d.) mm, film thickness: 1.5 m, 				coated with 5 % Phenyl and 95 % dimethylpolysiloxane (DB-5 				or equivalent)

Injection system

Injection volume	1 l

Detector

Temperatures

Column oven		Isothermal:	210°C

Injection port		250°C

Detector		280°C

Carrier gas		Helium, 10 ml/min

Retention times 		Cyprodinil: 	about 2.6 minutes

Sample preparation: Weigh (to the nearest 0.1 mg) sufficient sample (w mg) to contain about 50 mg cyprodinil into a volumetric flask (100 ml). Add acetone (about 80 ml) and place the flask in an ultrasonic bath for 5 min. Allow to cool to ambient temperature and fill to the mark with acetone. Mix thoroughly.

Calibration solution: Weigh (to the nearest 0.1 mg) sufficient sample (w mg) to contain about 50 mg cyprodinil into a volumetric flask (100 ml). Add acetone (about 80 ml) and place the flask in an ultrasonic bath for 5 min. Allow to cool to ambient temperature and fill to the mark with acetone. Mix thoroughly.

Typical Chromatogram 511/TC/M/3.1

Cyprodinil

2.3 Infrared. Prepare potassium bromide discs for the technical sample and reference substance. A typical KBr pellet will contain a sample prepared in the 0.2-0.4% by weight range. The spectra from the samples should not differ significantly from the spectrum of the reference substance.

Infrared Spectrum of Cyprodinil 511/TC/M/2.3

3 Cyprodinil

OUTLINE OF METHOD. Cyprodinil is determined by reversed phase high performance liquid chromatography using UV detection at 254 nm and external standard calibration.

3.1 Determination of Cyprodinil by reversed phase HPLC

- REAGENTS

- Cyprodinil reference standard of known content

Water HPLC grade

Acetonitrile HPLC grade

Trifluoroacetic acid analytical grade.

- Eluent: 1.0%v/v aqueous trifluoroacetic acid / acetonitrile 60 +40 (v/v)

Calibration solution. Weigh in duplicate (to the nearest 0.1 mg) about 100 mg of the cyprodinil reference standard (s mg) into separate volumetric flasks (100 ml). Add acetonitrile (about 40 ml) and place the flasks in an ultrasonic bath for 5 minutes. Allow the solutions to cool to ambient temperature and fill to the mark with acetonitrile. Mix thoroughly. Dilute 10.0 ml of the solutions with the eluent to a volume of 100.0 ml. Mix thoroughly (calibration solutions C1 and C2).

###### APPARATUS

High performance liquid chromatograph equipped with a detector suitable for operation at 254 nm (UV-detection) and an injection system capable to inject 10 µl.

Liquid chromatographic column stainless steel, 250 x 4.0 mm (i.d), packed with Nucleosil C18, 
5 µm, Macherey-Nagel or equivalent with the same selectivity.

Electronic integrator or data system

Ultrasonic bath

PROCEDURE

(a) Chromatographic conditions (typical)

Column temperature	ambient

Flow rate	1.2 ml/min

Detector wavelength	254 nm

Injection volume	10 µl

Retention time	approximately 9 min.

Eluent:	1.0%v/v aqueous trifluoroacetic acid / acetonitrile 60 +40 (v/v)

(b) Equilibration of the system. Pump sufficient mobile phase through the column to equilibrate the system. Inject 10 l portions of the calibration solution C1 and repeat the injections until retention times and peak areas deviate by less than  1 % from the mean for three successive injections.

- (c) Sample preparation. Prepare sample solutions in duplicate for each sample. Weigh (to the nearest 0.1 mg) sufficient sample (w mg) to contain about 100 mg cyprodinil into a volumetric flask (100 ml). Add acetonitrile (40 ml) and place the flasks in an ultrasonic bath for 5 minutes. Allow the solutions to cool to ambient temperature and fill to the mark with acetonitrile. Mix thoroughly. Dilute 10.0 ml of these solutions with the eluent to a volume of 100 ml. Mix thoroughly (sample solutions S1 and S2).

(d) Determination. Inject 10 l portions of the second calibration solution (C2) for three successive injections. The mean response factor for this solution should deviate by no more than 1% from those for the first calibration solution (C1) (see paragraph (b) Equilibration of the system), otherwise the calibration solutions should be prepared again.

Inject in duplicate 10 l portions of each sample solution (S1, S2, …, etc.) bracketing them by duplicate injections of the calibration solution (C1) using the following sequence:

- C1, C1, S1, S1, S2, S2, C1, C1…

Determine the peak area of cyprodinil.

- (e) Calculation

- Calculate the response factors (f) from the calibration solutions bracketing the injections of the sample solutions. Average the response factors of the calibration solutions preceding and following the sample solution injections. These must agree within 1 % of the average otherwise repeat the determination. Calculate the content of the sample solutions.

fi=s×PHs

- cyprodinil content =Hw×fw [g/kg] (M)

where:

fi	=	single response factor

f	=	average response factor

Hs	= peak area of cyprodinil in the calibration solution

Hw	= peak area of cyprodinil in the sample solution

s	= mass of cyprodinil reference standard in the calibration solution (mg)

w	= mass of sample taken (mg)

P	= purity of cyprodinil reference standard (g/kg)

###### Typical Chromatogram 511/TC/M/3.1

Cyprodinil

- CYPRODINIL EMULSIFIABLE CONCENTRATE

- 511/EC/M/-

1 Sampling.

Take at least 100 g.

2 Identity test

2.1 HPLC. As for cyprodinil technical 511/TC/M/2.1

2.2 GC. As for cyprodinil technical 511/TC/M/2.1

#### 3 Cyprodinil

Same approach as for cyprodinil technical 511/TC/M/3

#### 3.1 Determination of Cyprodinil by reversed phase HPLC

#### As for cyprodinil technical 511/TC/M/3.1 except

- (c) Sample preparation. Prepare sample solutions in duplicate for each sample. Weigh (to the nearest 0.1 mg) sufficient sample (w mg) to contain about 100 mg cyprodinil into a volumetric flask (100 ml). Suspend with 5 ml of water. Add acetonitrile (about 40 ml) and place the flasks in an ultrasonic bath for 5 minutes. Allow the solutions to cool to ambient temperature and fill to the mark with acetonitrile. Mix thoroughly. Dilute 10.0 ml of these solutions with the eluent to a volume of 100 ml. Mix thoroughly. Clear the suspension by filtration through a 0.45 m filter (sample solutions S1 and S2).

###### Typical Chromatogram 511/EC/M/3.1

Cyprodinil

- CYPRODINIL water dispersible granule

- 511/WG/M/-

1 Sampling

Take at least 100 g.

2 Identity test

2.1 HPLC. As for cyprodinil technical 511/TC/M/2.1

2.2 GC. As for cyprodinil technical 511/TC/M/2.1 except

Sample preparation: Weigh (to the nearest 0.1 mg) sufficient sample (w mg) to contain about 50 mg cyprodinil into a volumetric flask (100 ml). Add acetone (about 80 ml) and place the flask in an ultrasonic bath for 5 min. Allow to cool to ambient temperature and fill to the mark with acetone. Mix thoroughly and clear the suspension by filtration through a 0.45 m filter.

#### 3 Cyprodinil

Same approach as for cyprodinil technical 511/TC/M/3

#### 3.1 Determination of Cyprodinil by reversed phase HPLC
As for cyprodinil technical 511/TC/M/3.1 except

- (c) Sample preparation. Prepare sample solutions in duplicate for each sample. Weigh (to the nearest 0.1 mg) sufficient sample (w mg) to contain about 100 mg cyprodinil into a volumetric flask (100 ml). Suspend with 5 ml of water. Allow the suspension to stand for approximately 5 minutes. Add acetonitrile (about 40 ml) and place the flasks in an ultrasonic bath for 5 minutes. Allow the solutions to cool to ambient temperature and fill to the mark with acetonitrile. Mix thoroughly. Dilute 10.0 ml of these solutions with the eluent to a volume of 100 ml. Mix thoroughly. Clear the suspension by filtration through a 0.45 m filter (sample solutions S1 and S2).

###### Typical Chromatogram 511/WG/M/3.1

Cyprodinil

#### 4 Suspensibility

REAGENTS AND APPARATUS as for 511/TC/M/3.1 and MT 184.

add at:

APPARATUS

Sample filtering device with a membrane filtration unit compatible with organic solvents and a 0.45 μm pore diameter

PROCEDURE

(a) Preparation of suspension and determination of sedimentation. MT 184.

(b) Determination of Cyprodinil in the bottom 25 ml of suspension. After removal of the top 225 ml of suspension, transfer the remaining 25 ml quantitatively into a volumetric flask (100 ml) with acetonitrile (about 50 ml). Place the flask in an ultrasonic bath for 5 minutes. Allow to cool to ambient temperature and fill to the mark with acetonitrile. Mix thoroughly. Dilute 10.0 ml of the solution with the eluent to a volume of 100 ml. Mix thoroughly. Clear the suspension by filtration through a 0.45 μm filter prior to injection. Determine the mass (Q g) of Cyprodinil according to 511/WG/(M)/3.1, using a calibration solution with the appropriate final concentration of Cyprodinil standard.

(c) Calculation

Suspensibility = 111 ( c - Q )c %

where:

c	= mass of cyprodinil in the sample taken for the preparation of the suspension (g)

Q	= mass of cyprodinil in the bottom 25 ml of suspension (g)