Haloxyfop-P-Methyl

CIPAC Code No. 526

CIPAC Collaborative Study

STUDY TITLE

Analytical Method for Determination of Haloxyfop-P-Methyl

in Technical Material (TC) and EC Herbicide End-Use Products

by Normal Phase High Performance Liquid Chromatography

STUDY SPONSOR

Dow AgroSciences LLC

Crop Protection R&D Laboratories

Formulations Analytical and Product Chemistry Group

9330 Zionsville Road

Indianapolis, Indiana, USA 46268

AUTHOR

David Heim

TABLE OF CONTENTS

I.	LIST OF PARTICIPANTS	3

II.	INTRODUCTION	4

III.	SAMPLES AND STANDARDS	5

IV.	ANALYTICAL METHOD	5

1.	Scope	5

2.	Outline of Method	5

3.	Procedure	6

V.	REMARKS OF THE PARTICIPANTS	6

1.	Modifications to the Analytical Conditions	6

2.	Modifications to the Preparation of Calibration and Sample Solutions	6

3.	General Remarks on Performance of the Method	7

VI.	RESULTS AND DISCUSSION	8

1.	Method Performance	8

2.	Outliers and Stragglers	8

3.	Repeatability and Reproducibility	10

4.	Discussion	10

VII. CONCLUSIONS	10

TABLE I.	Results Summary by Laboratory  Haloxyfop-P-Methyl Technical	11

TABLE II.	Results Summary by Laboratory  Haloxyfop-P-Methyl EC Formulations	12

TABLE III.	Summary of Statistical Evaluation of the Collaborative Study Data	14

FIGURE 1.	Representative Chromatograms of Standard and Sample Solutions	15

FIGURE 2.	Graphical Presentation of Data  Haloxyfop-P-Methyl Technical	18

FIGURE 3.	Graphical Presentation of Data  Haloxyfop-P-Methyl EC Formulations	19

Appendix 	Supplemental Data: Collaborative Study Results and Statistical Analysis for Haloxyfop-methyl S-Enantiomer.	21

# I.	LIST OF PARTICIPANTS

| Jim Garvey | Pesticide Control Service Department of Agriculture and Food Ireland |
| --- | --- |
| Gerhard Genari | BASF Aktiengesellschaft Germany |
| David Heim | Dow AgroSciences LLC USA |
| Phil Jutsum | CEM Analytical Services UK |
| Helen Karasali George Balayiannis | Benaki Phytopathological Institute Chemical Analysis of Pesticides Greece |
| Luis Manso | Laboratorio Arbitral Agroalimentario Ministry of Agriculture Spain |
| Joe Moreland Karine Sevin | Dow AgroSciences S.A.S. France |
| Pia K. Nielsen | Cheminova A/S Analytical Development Laboratories Denmark |
| Olga Nováková Tat’ána Hulinová A. Št’astná | State Phytosanitary Administration Department of Chemical Laboratories Czech Republic |
| Bruno Patrian Astriid Bächli | Agroscope Changins-Wädenswil ACW Plant Protection Chemistry Switzerland |
| Andrew Plumb | Central Science Laboratory Food Science Group UK |
| Theo C. de Rijk | RIKILT – Institute of Food Safety The Netherlands |
| Juliana Schlosserova Frantisek Csicsay | Central Controlling and Testing Institute in Agriculture Department of Environmental Protection and Organic Farming Slovak Republic |
| Weili Shan Li Youshun Wang Xiaoli | Institute for the Control of Agrochemicals Ministry of Agriculture P. R. China |
| Tristan Speak | Dow AgroSciences LLC New Zealand |

# II.	INTRODUCTION

Common Name:	Haloxyfop-P-Methyl

CIPAC Number:	526

Chemical name	(+)-2-[4-[[3-chloro-5-(trifluoromethyl)-2-pyridinyl]oxy]phenoxy] propanoic acid methyl ester (CA; 72619-32-0)

Empirical formula	C16H13ClF3NO4

RMM	375.7

b.p.	> 280 C

v.p.	2.6 × 10-5 Pa at 20 C

Solubility	In water: 9.08 mg/l at 20 C

In most organic solvents:  >1000 g/l at 20 C

Description	Light amber, viscous liquid

Stability	Unstable in alkaline conditions

Formulations	Emulsifiable Concentrates (EC)

An analytical method for quantifying haloxyfop-P-methyl in technical material and EC formulations was developed and validated by Dow AgroSciences in 2007.  Following the successful pilot trial conducted in 2007 by members of the UK-PAC, participation in a full collaborative trial was requested in October 2007 as announced through CIPAC Information Sheet No. 276.  A protocol, data spreadsheet, MSDS, samples and standards were shipped to fifteen respondents. All fifteen participants completed the study (Section I) and provided results presented in this report.

# III.	SAMPLES AND STANDARDS

The following samples and standards were sent to the participants in December 2007.

| Sample Identification | Batch Identification | Dow AgroSciences Test Substance Number | Lot Number | Concentration /Purity |
| --- | --- | --- | --- | --- |
| Haloxyfop-P-Methyl Technical (TC) | TC1 | TSN106096 | VA08150101 | 969 g/kg 1 |
| Haloxyfop-P-Methyl Technical (TC) | TC2 | TSN106097 | VA09150101 | 969 g/kg 1 |
| EC1 Formulation | EC1 | TSN106098 | C1540-45-A | 36 g/kg 1 |
| EC2 Formulation | EC2 | TSN105767 | UF22150109 | 105 g/kg 1 |
| EC3 Formulation | EC3 | TSN106099 | 13-OCT-06-06 | 469 g/kg 1 |
| EC1 Formulation Blank |  | TSN106101 | C1540-45-B |  |
| EC2 Formulation Blank |  | TSN106100 | E2147-55 |  |
| EC3 Formulation Blank |  | TSN106102 | 195-95-A |  |
| Haloxyfop-P-Methyl Analytical Standard |  | TSN102596 | F0868-19 | 97.7% 2 1.4% 3 |
| Haloxyfop Methyl Racemic Standard |  | AGR229637 | ACPR 95-1 | 99.6% |

1	Nominal concentration of haloxyfop-P-methyl

2	Purity as R-enantiomer.  Chemical purity (R+S) = 99.1%, optical purity (R/R+S) = 98.6%.  Purity, 97.7% = 0.991  0.986  100%

3	Purity as S-enantiomer.  Chemical purity (R+S) = 99.1%, optical purity (S/R+S) = 1.4%.  Purity, 1.4% = 0.991  0.014  100%

# IV.	ANALYTICAL METHOD

## 1.	Scope

The method is applicable to the determination of haloxyfop-P-methyl in technical grade active ingredient and EC formulations.  The method has been validated over the range of 27 to 567 g/kg of haloxyfop-P-methyl in EC formulations.

## 2.	Outline of Method

An aliquot of the sample is diluted in a mixture of heptane and isopropanol and analyzed by normal phase HPLC using a Chiralcel OK column and ultraviolet (UV) detection at 280 nm.  Quantitation is by external standardization based on peak areas.

## 3.	Procedure

Two calibration standard solutions and two sample solutions of TC1, TC2, EC1, EC2, and EC3 were prepared on each day of analysis.  Samples were analyzed in duplicate on two days.  Prior to sample analyses on day 1, the mobile phase and formulation blanks were analyzed to evaluate interferences with haloxyfop-P-methyl and the racemic standard was analyzed to confirm appropriate resolution between enantiomers (R  1.5).  Prior to each day’s analyses, equilibration of the HPLC system was verified by comparison of peak areas between duplicate injections of a calibration standard solution.  Accurate preparation of the calibration standard solutions was confirmed by comparison of response factors.  Analytical standard, racemic standard, and TC samples were weighed into bottles and 50 mL of heptane – propan-2-ol, 80 + 20 (v/v), was added.  EC formulations and formulation blanks were weighed directly into 50-mL volumetric flasks and dissolved in heptane – propan-2-ol, 80 + 20 (v/v).  Representative chromatograms of standards and samples are presented in Figure 1.

# V.	REMARKS OF THE PARTICIPANTS

## 1.	Modifications to the Analytical Conditions

With a few exceptions noted below, all of the laboratories performed the analyses following the analytical conditions described in the study protocol.

|  | Column | Column Temp. | Mobile Phase | Flow Rate | Detection | Injection Volume |
| --- | --- | --- | --- | --- | --- | --- |
| Protocol Conditions | Chiralcel OK 4.6 × 250 mm 10 µm particles | 25C | Heptane – propan-2-ol – methanol, 85 + 5 + 10 (v/v); | 1.8 mL/min. | UV λ=280 nm | 10 µL |
| Lab 4 | Same | Same | Same | 1.1 mL/min. | Same | Same |
| Lab 6 | Same | Same | Same | 1.5 mL/min. | Same | Same |
| Lab 11 | Same | Same | Same | 1.5 mL/min.1 | Same | 20 µL |

1	Although column backpressure was normal at 1.8 mL/min., flow rate was reduced to avoid leakage from the detector during analysis.

## 2.	Modifications to the Preparation of Calibration and Sample Solutions

The amount of analytical standard and EC formulations weighed was reduced by the following laboratories.  The volume was correspondingly reduced so that the concentration of haloxyfop-P-methyl was consistent with a sample prepared following the method.

Lab 2:	The weight and volume for all calibration and sample solutions were halved.

Lab 8:	The weight and volume for EC1 formulation, day 2, replicate 2 was reduced by 1/5.

Lab 11:	The weight and volume for all calibration and sample solutions except EC3 formulation were halved.

Lab 14:	The weight and volume for all calibration and sample solutions were halved.

The amount of EC1 formulation blank provided to the laboratories was inadvertently less than required.  Since this occurred only for the EC1 formulation blank and would not compromise the results of the collaborative study, it was suggested to prepare the EC1 formulation blank solution by halving both the weight and volume for the solution.

## 3.	General Remarks on Performance of the Method

Lab 2:	Sample EC3 showed an interference peak occurring between the two enantiomer peaks.

Lab 4:	We utilized a new Chiralcel column off the shelf.  We observed an intense yellow color in the eluant prepared as in the method. The color faded somewhat in the course of the analyses, but remained clearly visible. We assume that the unspecific light absorption increased the noise of the UV detector but apparently did not affect both chromatographic resolution and quantitative results. In case other labs also made the same observation, we recommend to add a note accordingly.

The enantioselective HPLC-column seems to be quite slow in reaching equilibrium with the mobile phase. We achieved stable conditions only after overnight equilibration with the proper eluant and proper flow.

If these rules are kept in mind, the method is easy to use and reliable.

With the lower flow rate, we observed a better resolution of the additional small peak in the EC3 formulation eluting between R and S haloxyfop-methyl respectively. The blank EC3 formulation did not show this peak.

According to the instruction manual for Chiralcel OK columns, the maximum flow rate should not exceed the pressure upper's limit of 50 bar and the flow rate should not exceed 1.5ml/min. Therefore we used a flow rate of 1.1ml/min (pressure 42 bar). The retention times were longer in comparison with the protocol specification. We expanded the run time to 18min to compensate the lower flow rate. The chromatographic resolution still met the requirements of the instruction.

Lab 7:	Method could be implemented without any problems.

Lab 10:	First day was measured lower areas of standard then second day.  We think it was due to bad weighing of standard which is honey consistence.  When we were weighing standard second day we have to wash down standard many times with smaller amounts of solvent. We also think that 50 mL of solvent isn’t enough for quantitative transfer of this type consistence.  We wanted repeat analyses of technical compound and EC3 formulation 3rd day but we haven’t enough amount of analytical standard for duplicate weigh. We weighed standard for 3rd day once and we compared him with duplicate standards weighted for determination 2nd day. The areas of standards of second and third day correspond also areas of samples of first, second and third day correspond and for that reason I’m sure that the higher content of Haloxyfop-P-methyl is caused because of bad weight of standards.

I think we could avoid these problems with bad weighed standard if we did calibration curve before we started measuring. But for this procedure we haven’t enough amount of standard. We could see that standard is out of calibration curve.

I recommend bigger volume of solvent for quantitative transfer of standards and formulation from funnel or weighing bottle to volumetric flask especially for standards and technical compound which are honey consistence and keep at the surface. That can be reason why measured areas were lower.

Lab 15:	Retention time drifts backwards about half minutes in second day.

# VI.	RESULTS AND DISCUSSION

## 1.	Method Performance

Prior to this collaborative study and the pilot scale study, the analytical method was validated for appropriate accuracy, precision, linearity, and specificity.  During the collaborative study, specificity was evaluated to confirm that each column and system provided adequate separation of sample components.  The racemic standard and EC formulation blanks were used for this purpose, and each of the laboratories reported no interferences with haloxyfop-P-methyl.  Each laboratory also reported acceptable resolution of the S-enantiomer from haloxyfop-P-methyl (R 1.5) as required by the study protocol.

The response factors obtained for two consecutive injections of a calibration solution differed by less than 1.0% for all laboratories on each day indicating that the systems were properly equilibrated prior to each day’s analysis.  The response factor for the second calibration solution did not deviate by more than 1.7% from that for the first calibration solution suggesting that the laboratories correctly prepared the calibration solutions.  Although the response factors for the day 1 calibration solutions from laboratory 10 were within 0.5% of each other, comments and results from the laboratory suggest that both calibration solutions were invalid.

## 2.	Outliers and Stragglers

Fifteen sample sets were distributed, and results from each of the fifteen laboratories were obtained.  Laboratory data are presented in Tables I-II and are graphically presented in Figures 2-3.  The statistical evaluations for outlying results were completed according to “CIPAC Guidelines for Collaborative Study Procedures for Assessment of Performance of Analytical Methods” and ISO 5725-2:1994(E).  Outliers and stragglers were determined using Cochran’s test and Grubb’s test.  An outlying result was considered a straggler if the test statistic exceeded the critical value at =5% and was considered an outlier if the test statistic exceeded the critical value at =1%.  Cochran’s test outliers were removed first and the statistical evaluation was repeated.  If there were no Cochran’s test outliers after re-evaluation, Grubb’s test outliers were removed.  Following the CIPAC guidelines, data from only 3 of 15 laboratories could be removed without exceeding the 22% (2 of 9) limit.  Outliers and stragglers are presented in the following table and discussed below.

Summary of Outliers and Stragglers by Laboratory and Batch

| Test Sample | Technical |  |  |  | EC Formulation |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Batch | TC1 |  | TC2 |  | EC1 |  | EC2 |  | EC3 |  |
| Outlier Test | Cochran's | Grubb's | Cochran's | Grubb's | Cochran's | Grubb's | Cochran's | Grubb's | Cochran's | Grubb's |
| Lab No. |  |  |  |  |  |  |  |  |  |  |
| 1 |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  | ** |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |  |  |  |
| 6 |  |  |  |  |  | ** |  |  |  |  |
| 7 | * |  | ** |  |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |  |  |  |  |
| 10 | ** |  | ** | ** | * | * |  |  | ** |  |
| 11 |  |  | * |  |  |  |  |  |  |  |
| 12 |  |  |  |  |  |  | * |  |  |  |
| 13 |  |  |  |  | ** |  |  |  |  |  |
| 14 |  |  |  |  |  |  |  |  |  |  |
| 15 |  |  |  |  |  |  |  |  |  |  |

** Outlier	* Straggler

TC1.  The maximum variance (lab 10) was an outlier by Cochran’s test (Table Ia).  Following removal of lab 10, statistical evaluation was repeated, and the maximum variance (lab 4) was a straggler by Cochran’s test.  No other data were removed.

TC2.  The maximum variance and concentration (lab 10) were outliers by Cochran’s and Grubb’s test, respectively (Table Ib).  Following removal of lab 10, statistical evaluation was repeated, and the maximum variance (lab 7) was an outlier by Cochran’s test.  Following removal of lab 7, statistical evaluation was repeated, and the maximum variance (lab 2) was an outlier by Cochran’s test.  Following removal of lab 2, statistical evaluation was repeated, and the maximum variance (lab 11) was a straggler by Cochran’s test.

EC1.  The maximum variance (lab 13) was an outlier by Cochran’s test and the minimum concentration (lab 6) was a straggler by Grubbs’ test (Table IIa).  Following removal of lab 13, statistical evaluation was repeated, and the minimum concentration (lab 6) was an outlier by Grubb’s test.  Following removal of lab 6, statistical evaluation was repeated, and the maximum variance and concentration (lab 10) were stragglers by Cochran’s and Grubb’s tests, respectively.

EC2.  The maximum variance (lab 12) was a straggler by Cochran’s test (Table IIb).  No data were removed.

EC3.  The maximum variance (lab 10) was an outlier by Cochran’s test (Table IIc). Following removal of lab 10, statistical evaluation was repeated, and no other outliers or stragglers were found.

## 3.	Repeatability and Reproducibility

Statistical evaluation for repeatability and reproducibility is presented in Table III.  Before removal of outliers, the reproducibility relative standard deviations (RSDR) were 1.3%, 1.3%, 3.1%, 2.5%, and 1.7% for TC1, TC2, EC1, EC2, and EC3, respectively, and were each less than the Horwitz RSDR (Table IIIa).  Removal of outliers from TC1, TC2, EC1, and EC3 reduced the RSDR to 0.7%, 0.7%, 1.9%, and 1.4%, respectively (Table IIIb).  After outliers were removed, all RSDR were well below the Horwitz RSDR.  With the exception of laboratory 10 where the day 1 calibration solutions are suspected to be invalid, justification for rejection of outliers could not be identified.

Repeatability relative standard deviation (RSDr) values were less than the Horwitz RSDR before and after outlier removal, indicating a generally high level of repeatability within each laboratory.

## 4.	Discussion

The between-laboratory reproducibility, RSDR, was less than the Horwitz RSDR for each of the five batches of technical and EC formulations when all data were included in the statistical evaluation indicating that the method is reliable and robust.

# VII. CONCLUSIONS

Based on the relative standard deviation results, the method can provide acceptable reproducibility and repeatability according to the Horwitz equation.  Therefore, it is recommended that the method be accepted as provisional CIPAC method.

TABLE I.	Results Summary by Laboratory  Haloxyfop-P-Methyl Technical

A.  Technical Batch 1

|  | Mean Concentration (g/kg) |  |  |  |  | Standard Deviation | Mean Deviation |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Day 1 |  | Day 2 |  |  |  |  |
| Laboratory | Rep. 1 | Rep. 2 | Rep. 1 | Rep. 2 | Overall | (s) | (d) |
| 1 | 980 | 972 | 975 | 978 | 976.2 | 3.4 | 4.2 |
| 2 | 963 | 978 | 969 | 967 | 969.5 | 6.4 | -2.5 |
| 3 | 980 | 972 | 972 | 966 | 972.6 | 5.8 | 0.5 |
| 4 | 978 | 976 | 974 | 974 | 975.6 | 1.9 | 3.6 |
| 5 | 963 | 964 | 963 | 959 | 962.2 | 2.2 | -9.8 |
| 6 | 970 | 970 | 971 | 971 | 970.6 | 0.9 | -1.5 |
| 7 | 963 | 950 | 961 | 975 | 962.3 | 10.0*2 | -9.7 |
| 8 | 977 | 977 | 974 | 974 | 975.8 | 1.9 | 3.7 |
| 9 | 973 | 975 | 973 | 973 | 973.6 | 0.9 | 1.6 |
| 10 | 1033 | 1028 | 978 | 968 | 1001.9 | 33.6**1 | 29.9**1 |
| 11 | 970 | 959 | 953 | 971 | 963.1 | 8.8 | -8.9 |
| 12 | 962 | 965 | 963 | 960 | 962.7 | 2.3 | -9.3 |
| 13 | 978 | 973 | 976 | 976 | 975.7 | 2.4 | 3.7 |
| 14 | 976 | 969 | 980 | 970 | 974.0 | 5.2 | 1.9 |
| 15 | 969 | 965 | 959 | 965 | 964.6 | 4.1 | -7.4 |
| Mean |  |  |  |  | 972.0 |  |  |

* Straggler	1 Value(s) removed for re-evaluation of reproducibility and repeatability.

** Outlier	2 Value is a straggler after data from lab 10 is removed.

B.  Technical Batch 2

|  | Mean Concentration (g/kg) |  |  |  |  | Standard Deviation | Mean Deviation |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Day 1 |  | Day 2 |  |  |  |  |
| Laboratory | Rep. 1 | Rep. 2 | Rep. 1 | Rep. 2 | Overall | (s) | (d) |
| 1 | 970 | 980 | 972 | 974 | 973.9 | 4.0 | 1.1 |
| 2 | 975 | 978 | 948 | 973 | 968.6 | 13.7**1,3 | -4.2 |
| 3 | 970 | 972 | 969 | 969 | 969.9 | 1.2 | -2.9 |
| 4 | 980 | 981 | 979 | 980 | 980.3 | 0.9 | 7.5 |
| 5 | 964 | 966 | 962 | 965 | 964.3 | 1.9 | -8.5 |
| 6 | 969 | 969 | 972 | 970 | 970.1 | 1.5 | -2.7 |
| 7 | 993 | 993 | 960 | 970 | 979.0 | 16.3**1,2 | 6.2 |
| 8 | 976 | 978 | 973 | 965 | 973.0 | 5.8 | 0.2 |
| 9 | 972 | 971 | 975 | 971 | 972.2 | 1.9 | -0.6 |
| 10 | 1038 | 1008 | 976 | 974 | 998.9 | 30.5**1 | 26.1**1 |
| 11 | 969 | 988 | 974 | 967 | 974.4 | 9.6*4 | 1.6 |
| 12 | 962 | 968 | 970 | 965 | 966.1 | 3.2 | -6.7 |
| 13 | 977 | 976 | 972 | 972 | 974.2 | 2.9 | 1.4 |
| 14 | 975 | 967 | 963 | 956 | 965.3 | 7.7 | -7.5 |
| 15 | 961 | 958 | 968 | 963 | 962.1 | 4.2 | -10.7 |
| Mean |  |  |  |  | 972.8 |  |  |

* Straggler	1 Value(s) removed for re-evaluation of reproducibility and repeatability.

** Outlier	2 Value is an outlier after data from lab 10 is removed.

3 Value is an outlier after data from labs 10 and 7 are removed.

4 Value is a straggler after data from labs 10, 7, and 2 are removed.

TABLE II.	Results Summary by Laboratory  Haloxyfop-P-Methyl EC Formulations

A.  EC Formulation 1

|  | Mean Concentration (g/kg) |  |  |  |  | Standard Deviation | Mean Deviation |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Day 1 |  | Day 2 |  |  |  |  |
| Laboratory | Rep. 1 | Rep. 2 | Rep. 1 | Rep. 2 | Overall | (s) | (d) |
| 1 | 36.2 | 36.1 | 36.2 | 36.6 | 36.28 | 0.20 | 0.21 |
| 2 | 35.9 | 36.2 | 36.5 | 36.0 | 36.15 | 0.25 | 0.08 |
| 3 | 36.1 | 36.1 | 36.3 | 36.1 | 36.15 | 0.09 | 0.08 |
| 4 | 36.7 | 36.4 | 36.4 | 36.7 | 36.55 | 0.18 | 0.48 |
| 5 | 36.0 | 35.7 | 35.9 | 35.9 | 35.87 | 0.11 | -0.20 |
| 6 | 33.6 | 33.4 | 33.6 | 33.5 | 33.54 | 0.09 | -2.53**1,2 |
| 7 | 36.7 | 37.4 | 36.0 | 36.4 | 36.62 | 0.61 | 0.55 |
| 8 | 36.2 | 36.1 | 36.5 | 37.0 | 36.43 | 0.43 | 0.36 |
| 9 | 37.0 | 36.3 | 35.0 | 35.3 | 35.92 | 0.91 | -0.15 |
| 10 | 38.9 | 38.3 | 36.4 | 36.5 | 37.54 | 1.24*3 | 1.47*3 |
| 11 | 35.9 | 37.5 | 35.6 | 35.9 | 36.22 | 0.89 | 0.15 |
| 12 | 35.0 | 35.9 | 36.7 | 36.3 | 35.96 | 0.72 | -0.11 |
| 13 | 36.4 | 36.2 | 33.8 | 33.0 | 34.85 | 1.71**1 | -1.22 |
| 14 | 37.1 | 37.0 | 37.1 | 36.7 | 36.98 | 0.18 | 0.91 |
| 15 | 36.1 | 35.5 | 36.3 | 36.1 | 35.99 | 0.37 | -0.08 |
| Mean |  |  |  |  | 36.07 |  |  |

* Straggler	1 Value(s) removed for re-evaluation of reproducibility and repeatability.

** Outlier	2 Value is an outlier after data from lab 13 is removed.

3 Value is a straggler after data from labs 13 and 6 are removed.

B.  EC Formulation 2

|  | Mean Concentration (g/kg) |  |  |  |  | Standard Deviation | Mean Deviation |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Day 1 |  | Day 2 |  |  |  |  |
| Laboratory | Rep. 1 | Rep. 2 | Rep. 1 | Rep. 2 | Overall | (s) | (d) |
| 1 | 106 | 104 | 106 | 106 | 105.7 | 1.0 | 0.2 |
| 2 | 105 | 106 | 106 | 106 | 105.9 | 0.5 | 0.4 |
| 3 | 106 | 106 | 106 | 106 | 106.1 | 0.2 | 0.7 |
| 4 | 108 | 108 | 107 | 107 | 107.5 | 0.7 | 2.0 |
| 5 | 105 | 105 | 104 | 105 | 104.7 | 0.4 | -0.8 |
| 6 | 103 | 102 | 102 | 102 | 102.1 | 0.6 | -3.4 |
| 7 | 106 | 110 | 106 | 107 | 107.1 | 1.7 | 1.6 |
| 8 | 105 | 105 | 106 | 106 | 105.4 | 0.7 | 0.0 |
| 9 | 108 | 107 | 107 | 107 | 106.9 | 0.5 | 1.5 |
| 10 | 113 | 111 | 106 | 106 | 109.2 | 3.7 | 3.7 |
| 11 | 105 | 108 | 110 | 104 | 107.0 | 2.5 | 1.5 |
| 12 | 106 | 100 | 103 | 110 | 104.7 | 4.2* | -0.8 |
| 13 | 106 | 102 | 98.5 | 102 | 102.0 | 3.0 | -3.4 |
| 14 | 106 | 105 | 102 | 100 | 103.2 | 3.1 | -2.3 |
| 15 | 105 | 104 | 105 | 105 | 104.9 | 0.7 | -0.6 |
| Mean |  |  |  |  | 105.5 |  |  |

* Straggler

** Outlier

TABLE II (Cont.).	Results Summary by Laboratory  Haloxyfop-P-Methyl EC Formulations

C.  EC Formulation 3

|  | Mean Concentration (g/kg) |  |  |  |  | Standard Deviation | Mean Deviation |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Day 1 |  | Day 2 |  |  |  |  |
| Laboratory | Rep. 1 | Rep. 2 | Rep. 1 | Rep. 2 | Overall | (s) | (d) |
| 1 | 455 | 451 | 572 1 | 463 | 456.2 | 6.0 | -5.8 |
| 2 | 462 | 459 | 457 | 459 | 459.1 | 2.0 | -2.9 |
| 3 | 462 | 460 | 462 | 462 | 461.7 | 1.0 | -0.2 |
| 4 | 469 | 466 | 464 | 468 | 466.6 | 2.5 | 4.7 |
| 5 | 461 | 458 | 458 | 456 | 458.0 | 2.0 | -4.0 |
| 6 | 461 | 460 | 462 | 461 | 460.8 | 0.9 | -1.2 |
| 7 | 472 | 470 | 466 | 470 | 469.6 | 2.7 | 7.6 |
| 8 | 461 | 464 | 463 | 462 | 462.4 | 1.3 | 0.4 |
| 9 | 464 | 462 | 464 | 460 | 462.7 | 1.7 | 0.7 |
| 10 | 490 | 485 | 460 | 465 | 475.2 | 14.7**2 | 13.2 |
| 11 | 445 | 447 | 457 | 453 | 450.4 | 5.6 | -11.6 |
| 12 | 451 | 451 | 458 | 458 | 454.5 | 4.5 | -7.5 |
| 13 | 470 | 463 | 460 | 461 | 463.7 | 4.4 | 1.8 |
| 14 | 471 | 468 | 474 | 467 | 470.0 | 3.0 | 8.0 |
| 15 | 461 | 455 | 464 | 455 | 458.7 | 4.6 | -3.3 |
| Mean |  |  |  |  | 462.0 |  |  |

* Straggler	1 Rejected as an outlier.

** Outlier	2 Value(s) removed for re-evaluation of reproducibility and repeatability.

TABLE III.	Summary of Statistical Evaluation of the Collaborative Study Data

A.  No Test Results Removed

|  |  | Technical Batches |  | EC Formulations |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Batch Number |  | TC1 | TC2 |  | EC1 | EC2 | EC3 |
| Mean (x) |  | 972.0 | 972.8 |  | 36.07 | 105.5 | 462.0 |
| Standard Deviation (sx) |  | 9.9 | 8.9 |  | 0.91 | 2.0 | 6.5 |
| Laboratory Data Removed |  | None | None |  | None | None | None |
| Number of Labs (L) |  | 15 | 15 |  | 15 | 15 | 15 |
| Number of Analyses per Lab and Batch (n) |  | 4 | 4 |  | 4 | 4 | 4 |
| sr |  | 9.9 | 10.4 |  | 0.71 | 2.0 | 5.0 |
| sL |  | 8.6 | 7.2 |  | 0.84 | 1.7 | 5.9 |
| sR |  | 13.1 | 12.7 |  | 1.10 | 2.7 | 7.8 |
| r | sr*2.8 | 27.7 | 29.2 |  | 1.98 | 5.7 | 14.1 |
| R | sR*2.8 | 36.7 | 35.5 |  | 3.08 | 7.4 | 21.8 |
| RSDr | sr/x*100% | 1.0% | 1.1% |  | 2.0% | 1.9% | 1.1% |
| RSDR | sR/x*100% | 1.3% | 1.3% |  | 3.1% | 2.5% | 1.7% |
| RSDR(Horwitz) | 2(1-0.5log(Mean/100)) | 2.0% | 2.0% |  | 3.3% | 2.8% | 2.2% |

B.  Outliers Removed

|  |  | Technical Batches |  | EC Formulations |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Batch Number |  | TC1 | TC2 | EC1 | EC2 | EC3 |
| Mean (x) |  | 970.5 | 970.5 | 36.36 | 105.5 | 461.0 |
| Standard Deviation (sx) |  | 5.5 | 5.2 | 0.48 | 2.0 | 5.5 |
| Laboratory Data Removed |  | Labs 7, 10 | Labs 2, 7, 10 | Labs 6, 13 | None | Lab 10 |
| Number of Labs (L) |  | 13 | 12 | 13 | 15 | 14 |
| Number of Analyses per Lab and Batch (n) |  | 4 | 4 | 4 | 4 | 4 |
| sr |  | 4.2 | 4.6 | 0.59 | 2.0 | 3.4 |
| sL |  | 5.0 | 4.7 | 0.38 | 1.7 | 5.2 |
| sR |  | 6.6 | 6.6 | 0.70 | 2.7 | 6.3 |
| r | sr*2.8 | 11.8 | 12.8 | 1.66 | 5.7 | 9.6 |
| R | sR*2.8 | 18.4 | 18.4 | 1.96 | 7.4 | 17.6 |
| RSDr | sr/x*100% | 0.4% | 0.5% | 1.6% | 1.9% | 0.7% |
| RSDR | sR/x*100% | 0.7% | 0.7% | 1.9% | 2.5% | 1.4% |
| RSDR(Horwitz) | 2(1-0.5log(Mean/100)) | 2.0% | 2.0% | 3.3% | 2.8% | 2.2% |

FIGURE 1.	Representative Chromatograms of Standard and Sample Solutions

A.   Representative Chromatogram for Haloxyfop Methyl Racemic Standard

B.   Representative Chromatogram for Haloxyfop-P-Methyl Calibration Standard Solution

FIGURE 1 (Cont.).	Representative Chromatograms of Standard and Sample Solutions

C.   Representative Chromatogram for a TC Sample

D.   Representative Chromatogram for EC1 Formulation

FIGURE 1 (Cont.).	Representative Chromatograms of Standard and Sample Solutions

E.   Representative Chromatogram for EC2 Formulation

F.   Representative Chromatogram for EC3 Formulation

FIGURE 2.	Graphical Presentation of Data  Haloxyfop-P-Methyl Technical

| Technical Batch 1 (TC1) |  |
| --- | --- |
| Technical Batch 2 (TC2) |  |

FIGURE 3.	Graphical Presentation of Data  Haloxyfop-P-Methyl EC Formulations

| EC1 Formulation |  |
| --- | --- |
| EC2 Formulation |  |

FIGURE 3 (Cont.).	Graphical Presentation of Data  Haloxyfop-P-Methyl EC Formulations

| EC3 Formulation |  |
| --- | --- |

Appendix 	Supplemental Data: Collaborative Study Results and Statistical Analysis for Haloxyfop-methyl S-Enantiomer.

Following presentation of the pilot trial results in June 2007, CIPAC requested collection of data on the S-enantiomer.  Collection of the S-enantiomer data was accomplished by requesting the laboratories in the collaborative study to integrate the S-enantiomer peak in all standard and samples solutions.   The haloxyfop-P-methyl analytical standard, Dow AgroSciences test substance number TSN102596, for which the content of S-enantiomer was known, was used as the calibration standard.  The results were tabulated and statistical evaluation was performed in the same manner as for the R-enantiomer.  While the analytical method was validated for the R-enantiomer including resolution from the S-enantiomer, the method was not validated for quantifying the S-enantiomer. Therefore, the data on the following pages is considered supplemental and should be used with caution.

TABLE IV.	Results Summary by Laboratory  S-Enantiomer in Haloxfop-P-Methyl Technical

A.  Technical Batch 1

|  | Mean Concentration (g/kg) |  |  |  |  | Standard Deviation | Mean Deviation |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Day 1 |  | Day 2 |  |  |  |  |
| Laboratory | Rep. 1 | Rep. 2 | Rep. 1 | Rep. 2 | Overall | (s) | (d) |
| 1 | 7.2 | 7.2 | 7.1 | 7.2 | 7.2 | 0.06 | -0.55 |
| 2 | 7.5 | 7.7 | 7.7 | 7.6 | 7.6 | 0.10 | -0.08 |
| 3 | 7.5 | 7.5 | 7.5 | 7.5 | 7.5 | 0.01 | -0.22 |
| 4 | 7.7 | 7.7 | 7.7 | 7.7 | 7.7 | 0.01 | 0.00 |
| 5 | 6.9 | 7.5 | 7.4 | 6.9 | 7.2 | 0.32 | -0.53 |
| 6 | 7.6 | 7.6 | 7.6 | 7.6 | 7.6 | 0.02 | -0.12 |
| 7 | 7.9 | 9.1 | 8.8 | 7.2 | 8.2 | 0.87 | 0.53 |
| 8 | 8.1 | 8.5 | 8.3 | 8.4 | 8.3 | 0.14 | 0.61 |
| 9 | 7.1 | 7.1 | 7.1 | 7.1 | 7.1 | 0.04 | -0.59 |
| 10 | 9.4 | 9.7 | 8.1 | 8.3 | 8.9 | 0.79 | 1.17 |
| 11 | 7.5 | 7.4 | 7.4 | 7.4 | 7.4 | 0.02 | -0.26 |
| 12 | 11.7 | 7.8 | 7.5 | 7.5 | 8.6 | 2.03**1 | 0.94 |
| 13 | 7.8 | 7.6 | 7.2 | 7.3 | 7.4 | 0.28 | -0.26 |
| 14 | 10.4 | 9.3 | 9.5 | 9.5 | 9.7 | 0.47 | 1.97 |
| 15 | 5.6 | 5.8 | 3.8 | 5.1 | 5.1 | 0.90*2 | -2.61 |
| Mean |  |  |  |  | 7.7 |  |  |

* Straggler	1 Value(s) removed for re-evaluation of reproducibility and repeatability.

** Outlier	2 Value is a straggler after data from lab 12 are removed.

B.  Technical Batch 2

|  | Mean Concentration (g/kg) |  |  |  |  | Standard Deviation | Mean Deviation |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Day 1 |  | Day 2 |  |  |  |  |
| Laboratory | Rep. 1 | Rep. 2 | Rep. 1 | Rep. 2 | Overall | (s) | (d) |
| 1 | 7.5 | 7.6 | 7.4 | 7.2 | 7.5 | 0.17 | -0.68 |
| 2 | 7.5 | 8.0 | 7.8 | 8.0 | 7.8 | 0.24 | -0.29 |
| 3 | 7.9 | 7.9 | 7.8 | 7.9 | 7.9 | 0.04 | -0.25 |
| 4 | 8.1 | 8.0 | 8.1 | 8.1 | 8.1 | 0.04 | -0.03 |
| 5 | 7.2 | 7.3 | 7.6 | 8.3 | 7.6 | 0.52 | -0.51 |
| 6 | 7.9 | 7.9 | 8.0 | 8.0 | 7.9 | 0.03 | -0.18 |
| 7 | 10.0 | 10.0 | 9.2 | 9.7 | 9.7 | 0.37 | 1.60 |
| 8 | 8.8 | 8.6 | 8.6 | 8.8 | 8.7 | 0.10 | 0.56 |
| 9 | 7.6 | 7.6 | 7.6 | 7.5 | 7.5 | 0.03 | -0.58 |
| 10 | 9.7 | 9.4 | 8.4 | 8.7 | 9.0 | 0.60*2 | 0.92 |
| 11 | 7.9 | 8.3 | 7.8 | 7.7 | 7.9 | 0.27 | -0.21 |
| 12 | 8.8 | 8.1 | 7.9 | 8.0 | 8.2 | 0.40 | 0.08 |
| 13 | 7.8 | 8.0 | 7.7 | 7.8 | 7.8 | 0.14 | -0.31 |
| 14 | 9.5 | 9.6 | 9.7 | 9.9 | 9.7 | 0.18 | 1.56 |
| 15 | 5.6 | 6.5 | 5.5 | 8.1 | 6.4 | 1.20**1 | -1.70 |
| Mean |  |  |  |  | 8.1 |  |  |

* Straggler	1 Value(s) removed for re-evaluation of reproducibility and repeatability.

** Outlier	2 Value is an outlier after data from lab 15 are removed.

TABLE V.	Results Summary by Laboratory  S-Enantiomer in Haloxyfop-P-Methyl EC Formulations

A.  EC Formulation 1

|  | Mean Concentration (g/kg) |  |  |  |  | Standard Deviation | Mean Deviation |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Day 1 |  | Day 2 |  |  |  |  |
| Laboratory | Rep. 1 | Rep. 2 | Rep. 1 | Rep. 2 | Overall | (s) | (d) |
| 1 | 0.28 | 0.28 | 0.28 | 0.28 | 0.28 | 0.002 | -0.014 |
| 2 | 0.28 | 0.30 | 0.30 | 0.30 | 0.29 | 0.009 | 0.000 |
| 3 | 0.29 | 0.29 | 0.29 | 0.29 | 0.29 | 0.002 | -0.003 |
| 4 | 0.30 | 0.30 | 0.30 | 0.30 | 0.30 | 0.003 | 0.005 |
| 5 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | 0.003 | -0.024 |
| 6 | 0.27 | 0.27 | 0.27 | 0.27 | 0.27 | 0.003 | -0.024 |
| 7 | 0.37 | 0.37 | 0.28 | 0.35 | 0.34 | 0.042**1,2 | 0.049* |
| 8 | 0.29 | 0.28 | 0.26 | 0.32 | 0.29 | 0.028*3 | -0.006 |
| 9 | 0.27 | 0.27 | 0.27 | 0.28 | 0.27 | 0.003 | -0.023 |
| 10 | 0.29 | 0.30 | 0.30 | 0.32 | 0.30 | 0.011 | 0.008 |
| 11 | 0.29 | 0.32 | 0.29 | 0.29 | 0.30 | 0.013 | 0.003 |
| 12 | 0.27 | 0.28 | 0.29 | 0.29 | 0.28 | 0.010 | -0.010 |
| 13 | 0.30 | 0.29 | 0.27 | 0.26 | 0.28 | 0.017 | -0.014 |
| 14 | 0.36 | 0.37 | 0.40 | 0.45 | 0.39 | 0.043* | 0.101**1 |
| 15 | 0.22 | 0.27 | 0.23 | 0.27 | 0.25 | 0.026 | -0.048 |
| Mean |  |  |  |  | 0.29 |  |  |

* Straggler	1 Value(s) removed for re-evaluation of reproducibility and repeatability.

** Outlier	2 Value is an outlier after data from lab 14 are removed.

3 Value is a straggler after data from labs 14 and 7 are removed.

B.  EC Formulation 2

|  | Mean Concentration (g/kg) |  |  |  |  | Standard Deviation | Mean Deviation |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Day 1 |  | Day 2 |  |  |  |  |
| Laboratory | Rep. 1 | Rep. 2 | Rep. 1 | Rep. 2 | Overall | (s) | (d) |
| 1 | 0.93 | 0.89 | 0.94 | 0.95 | 0.93 | 0.03 | -0.02 |
| 2 | 0.90 | 0.98 | 0.99 | 0.99 | 0.96 | 0.04 | 0.01 |
| 3 | 1.01 | 0.96 | 0.97 | 0.96 | 0.97 | 0.02 | 0.02 |
| 4 | 1.00 | 1.00 | 1.01 | 1.00 | 1.00 | 0.00 | 0.05 |
| 5 | 0.91 | 0.87 | 0.90 | 0.89 | 0.89 | 0.02 | -0.06 |
| 6 | 0.94 | 0.95 | 0.94 | 0.95 | 0.95 | 0.01 | -0.01 |
| 7 | 1.07 | 1.08 | 0.99 | 0.83 | 0.99 | 0.12 | 0.04 |
| 8 | 0.67 | 0.71 | 0.84 | 0.94 | 0.79 | 0.12 | -0.16* |
| 9 | 0.97 | 1.01 | 0.99 | 1.00 | 0.99 | 0.02 | 0.04 |
| 10 | 1.19 | 1.03 | 0.95 | 0.95 | 1.03 | 0.12 | 0.08 |
| 11 | 0.99 | 0.99 | 1.09 | 0.95 | 1.01 | 0.06 | 0.05 |
| 12 | 1.16 | 0.99 | 0.90 | 0.97 | 1.01 | 0.11 | 0.05 |
| 13 | 1.02 | 0.94 | 0.90 | 0.92 | 0.95 | 0.05 | -0.01 |
| 14 | 0.89 | 0.76 | 0.84 | 1.07 | 0.89 | 0.13 | -0.06 |
| 15 | 0.92 | 0.90 | 0.95 | 0.95 | 0.93 | 0.02 | -0.02 |
| Mean |  |  |  |  | 0.95 |  |  |

* Straggler

** Outlier

TABLE V (Cont.).	Results Summary by Laboratory  S-Enantiomer in Haloxyfop-P-Methyl EC Formulations

C.  EC Formulation 3

|  | Mean Concentration (g/kg) |  |  |  |  | Standard Deviation | Mean Deviation |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Day 1 |  | Day 2 |  |  |  |  |
| Laboratory | Rep. 1 | Rep. 2 | Rep. 1 | Rep. 2 | Overall | (s) | (d) |
| 1 | 2.5 | 2.2 | 3.0 | 2.5 | 2.5 | 0.34* | -2.33 |
| 2 | 5.5 | 5.3 | 5.4 | 5.6 | 5.4 | 0.11 | 0.57 |
| 3 | 5.6 | 5.6 | 5.7 | 5.7 | 5.6 | 0.06 | 0.78 |
| 4 | 3.8 | 3.8 | 3.9 | 3.9 | 3.9 | 0.02 | -1.02 |
| 5 | 4.9 | 4.7 | 4.8 | 4.7 | 4.8 | 0.06 | -0.10 |
| 6 | 5.3 | 5.4 | 5.6 | 5.7 | 5.5 | 0.16 | 0.64 |
| 7 | 6.1 | 6.3 | 5.9 | 5.9 | 6.0 | 0.19 | 1.17 |
| 8 | 4.5 | 4.5 | 4.6 | 4.7 | 4.6 | 0.13 | -0.29 |
| 9 | 5.6 | 5.6 | 5.6 | 5.6 | 5.6 | 0.03 | 0.75 |
| 10 | 5.6 | 5.1 | 5.0 | 4.9 | 5.1 | 0.30 | 0.25 |
| 11 | 4.2 | 4.1 | 4.2 | 4.3 | 4.2 | 0.09 | -0.67 |
| 12 | 6.1 | 9.0 | 5.2 | 4.6 | 6.2 | 1.97**1 | 1.36 |
| 13 | 2.8 | 2.8 | 2.8 | 2.8 | 2.8 | 0.01 | -2.04 |
| 14 | 5.2 | 5.3 | 5.4 | 5.6 | 5.4 | 0.17 | 0.54 |
| 15 | 5.0 | 5.4 | 5.2 | 5.3 | 5.2 | 0.15 | 0.36 |
| Mean |  |  |  |  | 4.9 |  |  |

* Straggler	1 Value(s) removed for re-evaluation of reproducibility and repeatability.

** Outlier	2 Value is a straggler after data from lab 12 are removed.

TABLE VI.	Summary of Statistical Evaluation of the Collaborative Study Data  S-Enantiomer

A.  No Test Results Removed

|  |  | Technical Batches |  | EC Formulations |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Batch Number |  | TC1 | TC2 |  | EC1 | EC2 | EC3 |
| Mean (x) |  | 7.7 | 8.1 |  | 0.29 | 0.95 | 4.9 |
| Standard Deviation (sx) |  | 1.03 | 0.86 |  | 0.035 | 0.061 | 1.09 |
| Laboratory Data Removed |  | None | None |  | None | None | None |
| Number of Labs (L) |  | 15 | 15 |  | 15 | 15 | 15 |
| Number of Analyses per Lab and Batch (n) |  | 4 | 4 |  | 4 | 4 | 4 |
| sr |  | 0.67 | 0.42 |  | 0.020 | 0.074 | 0.53 |
| sL |  | 0.97 | 0.84 |  | 0.034 | 0.049 | 1.06 |
| sR |  | 1.18 | 0.93 |  | 0.039 | 0.089 | 1.19 |
| r | sr*2.8 | 1.88 | 1.16 |  | 0.055 | 0.208 | 1.49 |
| R | sR*2.8 | 3.30 | 2.61 |  | 0.109 | 0.249 | 3.33 |
| RSDr | sr/x*100% | 8.7% | 5.1% |  | 6.7% | 7.8% | 10.9% |
| RSDR | sR/x*100% | 15.3% | 11.5% |  | 13.2% | 9.3% | 24.4% |
| RSDR(Horwitz) | 2(1-0.5log(Mean/100)) | 4.2% | 4.1% |  | 6.8% | 5.7% | 4.5% |

B.  Outliers Removed

|  |  | Technical Batches |  | EC Formulations |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Batch Number |  | TC1 | TC2 | EC1 | EC2 | EC3 |
| Mean (x) |  | 7.6 | 8.2 | 0.28 | 0.95 | 4.8 |
| Standard Deviation (sx) |  | 1.03 | 0.75 | 0.015 | 0.061 | 1.07 |
| Laboratory Data Removed |  | Labs 7, 10 | Labs 2, 7, 10 | Labs 6, 13 | None | Lab 10 |
| Number of Labs (L) |  | 14 | 14 | 13 | 15 | 14 |
| Number of Analyses per Lab and Batch (n) |  | 4 | 4 | 4 | 4 | 4 |
| sr |  | 0.43 | 0.29 | 0.013 | 0.074 | 0.16 |
| sL |  | 1.01 | 0.74 | 0.014 | 0.049 | 1.06 |
| sR |  | 1.10 | 0.79 | 0.019 | 0.089 | 1.08 |
| r | sr*2.8 | 1.22 | 0.80 | 0.037 | 0.208 | 0.46 |
| R | sR*2.8 | 3.07 | 2.21 | 0.054 | 0.249 | 3.01 |
| RSDr | sr/x*100% | 5.7% | 3.5% | 4.7% | 7.8% | 3.4% |
| RSDR | sR/x*100% | 14.4% | 9.6% | 6.8% | 9.3% | 22.6% |
| RSDR(Horwitz) | 2(1-0.5log(Mean/100)) | 4.2% | 4.1% | 6.8% | 5.7% | 4.5% |

FIGURE 4.	Graphical Presentation of Data  S-Enantiomer in Haloxyfop-P-Methyl Technical

| Technical Batch 1 (TC1) |  |
| --- | --- |
| Technical Batch 2 (TC2) |  |

FIGURE 5.	Graphical Presentation of Data  S-Enantiomer in Haloxyfop-P-Methyl EC Formulations

| EC1 Formulation |  |
| --- | --- |
| EC2 Formulation |  |

FIGURE 5 (Cont.).	Graphical Presentation of Data  S-Enantiomer in Haloxyfop-P-Methyl EC Formulations

| EC3 Formulation |  |
| --- | --- |