653. Cyazofamid

HPLC method 4773m

CIPAC Small Scale Collaborative Trial

##### Frédéric Joris,

ISK Biosciences Europe N.V.
B-1831 Diegem

- Belgium

##### January 2011

### cyAzofamid

653

- ISO common name:	Cyazofamid

Chemical name:   	4-chloro-2cyano-N,N-dimethyl-5-P-tolylimidazole -1- sulfonamide (IUPAC)

4-chloro-2cyano-N,N-dimethyl-5-(4-methylphenyl)-1H- imidazole- 1-sulfonamide (CA)

CAS-Number:  	120116-88-3

RMM: 			324.8

Empirical formula:	C13H13ClN4O2S

m.p.	152.7 °C

V.p.	<1.33x10 -5 Pa at 35 °C

Solubility	In water at 20 °C: 0.107-0.121 mg/l at pH 5, 7 and 9

acetone: 41.92 g/l at 20 °C

dichloromethane: 101.84 g/l at 20 °C

ethyl acetate: 15.63 g/l at 20 °C

hexane: 0.03 g/l at 20 °C

methanol: 1.54 g/l at 20 °C

n-octanol: 0.25 g/l at 20 °C

toluene: 5.28 g/l at 20 °C

acetonitrile: 29.42 g/l at 20 °C

propan-2-ol: 0.39 g/l at 20 °C

Stability	Keep frozen (< -18 °C) when not in use and avoid exposure to light.

Hydrolysis	DT50 at 20 °C: 24.6 days (pH 4), 27.2 days (pH 5), 24.8 days (pH 7) and 24.8 days (pH 9)

Description	White powder

Formulation	Suspension concentrates

### Cyazofamid technical

653/TC/M/-

1 Sampling. Take at least 20 g.

## 2 Identity test

2.1 HPLC. Use the reversed phase HPLC method described below. The relative retention time of the cyazofamid peak in the sample solution should not deviate by more than 2% from that of the calibration solution. The UV spectrum measured from this peak should match that obtained from the calibration substance.

- 3 Cyazofamid

OUTLINE OF METHOD. Cyazofamid is determined by reversed phase high performance liquid chromatography using UV detection at 280 nm and external standard calibration.

- REAGENTS

- Cyazofamid reference standard with known content

Water HPLC grade

Acetic acid analytical reagent grade

Acetonitrile HPLC grade

Methanol HPLC grade

- Mobile phase: water (pH 4 with acetic acid) – acetonitrile – methanol (43-32-25 v/v)

Solvent: acetonitrile

Calibration solutions. Weigh in duplicate (to the nearest 0.1 mg) 85 mg of cyazofamid reference standard (s mg) into separate volumetric flasks (50 ml). Add acetonitrile (about     40 ml) and sonicate for 5 minutes until complete dissolution. Allow the solutions to cool to ambient temperature and fill to the mark with acetonitrile (calibration solutions CA and CB). Mix well.

###### APPARATUS

High performance liquid chromatograph equipped with a detector suitable for operation at 280 nm (UV-detection) and an injection system capable of injecting 5 µl

Liquid chromatographic column stainless steel, 250 x 4.6 mm (i.d), packed with C18, 
5 µm, Phenomenex SphereClone ODS-2 or equivalent with the same selectivity

Electronic integrator or data system

Ultrasonic bath

pHmeter

PROCEDURE

(a) Chromatographic conditions (typical)

Autosampler temperature	  10°C (Integrated chiller)

Column temperature		30°C

Flow rate		1.0 ml/min

Detector wavelength		280 nm

Injection volume		5 µl

Retention time		approximately 27.7 min

- Mobile phase	water (pH 4 with acetic acid) – acetonitrile – methanol (43-32-25 v/v)

- Elution				isocratic

(b) Equilibration of the system. Pump sufficient mobile phase through the column to equilibrate the system. Inject 5 l portions of the calibration solution CA and repeat the injections until retention times and peak areas deviate by less than  0.5 % from the mean for three successive injections.

- (c) Sample preparation. Prepare sample solutions in duplicate for each sample. Weigh (to the nearest 0.1 mg) sufficient sample (w mg) to contain about 85 mg cyazofamid into a volumetric flask (50 ml). Add acetonitrile (about 40 ml) and sonicate for 5 minutes until complete dissolution. Allow the solutions to cool to ambient temperature and fill to the mark with acetonitrile (sample solutions S1 and S2). Mix thoroughly.

(d) Determination. Inject 5 l portions of the second calibration solution (CB) for two successive injections. The mean response factor for this solution should deviate by no more than 1% from that for the first calibration solution (CA) (see paragraph (b) Equilibration of the system), otherwise the calibration solutions should be prepared again.

Inject in duplicate 5 l portions of each sample solution (S1, S2, …,etc.) bracketing them by single injections of calibration solutions (CA and CB) using the following sequence:

- CA, S1, S1, S2, S2, CB, S3, S3, S4, S4, CA…

(e) 	Calculation. Determine the peak area of cyazofamid and calculate the mean value of response factors from the calibration solutions bracketing the injections of the sample solutions and use this value for calculating the cyazofamid content of the bracketed sample solutions. The cyazofamid content is the mean value of two sample solutions.

- fi=s×PHs

- cyazofamid content =f×Hww [g/kg] (M)

where:

fi	= individual response factor

f	= mean response factor

Hs	= peak area of cyazofamid in the calibration solution

Hw	= peak area of cyazofamid in the sample solution

s	= mass of cyazofamid reference standard in the calibration solution (mg)

w   = mass of sample taken (mg)

P	= purity of cyazofamid reference standard (g/kg)

###### Typical Chromatograms 653/TC/M/3

Calibration solution of CYAZOFAMID

Sample solution of CYAZOFAMID TECHNICAL

Solution of solvent (acetonitrile)

CYAZOFAMID SUSPENSION CONCENTRATES

- 653/SC/M/-

1 Sampling. Take at least 100 g.

2 Identity test

2.1 HPLC. As for cyazofamid technical 653/TC/M/2.1

#### 3 Cyazofamid

#### As for cyazofamid technical 653/TC/M/3 except

- (c) Sample preparation. Prepare sample solutions in duplicate for each sample. Weigh (to the nearest 0.1 mg) sufficient sample (w mg) to contain about 85 mg cyazofamid into a volumetric flask (50 ml). Add acetonitrile (about 40 ml) and sonicate for 15 minutes until complete dissolution. Allow the solutions to cool to ambient temperature and fill to the mark with acetonitrile (sample solutions S1 and S2). Mix thoroughly. Filter an aliquot of each prepared solution through a 0.45 µm PTFE filter prior to analysis.