595. Flazasulfuron

HPLC method 4791m

CIPAC Small Scale Collaborative Trial

##### Frédéric Joris,

ISK Biosciences Europe N.V.
B-1831 Diegem

- Belgium

##### January 2011

### FLazasULFURON

595

- ISO common name:	Flazasulfuron

Chemical name:   	1-(4,6-dimethoxypyrimidin-2-yl)-3-(3-trifluoromethyl-2-pyridylsulfonyl)urea (IUPAC)

CAS-Number:  	104040-78-0

RMM: 			407.36

Empirical formula:	C13H12F3N5O5S

m.p.	180 °C

V.p.	< 1.33 x 10-5 Pa at 25°C, 35°C, 45 °C

Solubility	In water at 25 °C: 2.1 g/l at pH 7, 0.027 g/l at pH 5 and not stable         at pH 9

In n-hexane : 0.5 mg/l at 25°C

In toluene: 0.56 g/l at 25°C

In dichloromethane: 22.1 g/l at 25°C

In methanol: 4.2 g/l at 25°C

In acetone: 22.7 g/l at 25°C

In ethyl acetate: 6.9 g/l at 25°C

In n-octanol: 0.20 g/l at 25°C

In acetonitrile: 8.7 g/l at 25°C

Stability	Keep frozen (< -18 °C) when not in use and avoid exposure to light.

Hydrolysis	DT50 at 22 °C: 16.6 days (pH 7), 13.1 days (pH 9) and 17.4 hours (pH 4)

Description	Form : White powder

Formulation	Water dispersible granules

### FLAZASULFURON technical

595/TC/M/-

1 Sampling. Take at least 20 g.

## 2 Identity test

2.1 HPLC. Use the reversed phase HPLC method 3 described below. The relative retention time of the flazasulfuron peak in the sample solution should not deviate by more than 2% from that of the calibration solution. The UV spectrum measured from this peak should match that obtained from the calibration substance.

- 3 Flazasulfuron

OUTLINE OF METHOD. Flazasulfuron is determined by reversed phase high performance liquid chromatography using UV detection at 230 nm and external standard calibration.

- REAGENTS

- Flazasulfuron reference standard of known content

Water HPLC grade

Acetonitrile HPLC grade

Acetic acid, 99.5%, analytical reagent grade

- Mobile phase: water (0.05% acetic acid):acetonitrile (45:55, v/v)

Solvent: acetonitrile

Calibration solutions. Weigh in duplicate (to the nearest 0.1 mg) 40 mg of flazasulfuron reference standard (s mg) into separate volumetric flasks (100 ml). Add acetonitrile (about   75 ml) and sonicate for 5 minutes until complete dissolution. Allow the solutions to cool to ambient temperature and fill to the mark with acetonitrile (calibration solutions CA and CB). Mix well.

###### APPARATUS

High performance liquid chromatograph equipped with a detector suitable for operation at 230 nm (UV-detection) and an injection system capable of injecting 10 µl

Liquid chromatographic column stainless steel, 250 x 4.6 mm (i.d), packed with C18, 
5 µm, Zorbax Eclipse XDB or equivalent with the same selectivity

Electronic integrator or data system

Ultrasonic bath

PROCEDURE

(a) Chromatographic conditions (typical)

Column temperature		Ambient

Flow rate		1.0 ml/min

Detector wavelength		230 nm

Injection volume		10 µl

Retention time		approximately 6.5 min

- Mobile phase	 		water (0.05% acetic acid):acetonitrile (45:55, v/v)

- Elution					isocratic

(b) Equilibration of the system. Pump sufficient mobile phase through the column to equilibrate the system. Inject 10 l portions of the calibration solution CA and repeat the injections until retention times and peak areas deviate by less than  0.5 % from the mean for three successive injections.

- (c) Sample preparation. Prepare sample solutions in duplicate for each sample. Weigh (to the nearest 0.1 mg) sufficient sample (w mg) to contain about 40 mg flazasulfuron into a volumetric flask (100 ml). Add acetonitrile (about 75 ml) and sonicate for 5 minutes until complete dissolution. Allow the solutions to cool to ambient temperature and fill to the mark with acetonitrile (sample solutions S1 and S2). Mix thoroughly. Filter an aliquot of each prepared solution through a 0.45 µm PTFE filter prior to analysis.

(d) Determination. Inject 10 l portions of the second calibration solution (CB) for two successive injections. The mean response factor for this solution should deviate by no more than 1% from that for the first calibration solution (CA) (see paragraph (b) Equilibration of the system), otherwise the calibration solutions should be prepared again.

Inject in duplicate 10 l portions of each sample solution (S1, S2, …,etc.) bracketing them by single injections of calibration solutions (CA and CB) using the following sequence:

- CA, S1, S1, S2, S2, CB, S3, S3, S4, S4, CA…

(e) 	Calculation. Determine the peak area of flazasulfuron and calculate the mean value of response factors from the calibration solutions bracketing the injections of the sample solutions and use this value for calculating the flazasulfuron content of the bracketed sample solutions. The flazasulfuron content is the mean value of two sample solutions.

- fi=s×PHs

- flazasulfuron content =f×Hww [g/kg] (M)

where:

fi	= individual response factor

f	= mean response factor

Hs	= peak area of flazasulfuron in the calibration solution

Hw	= peak area of flazasulfuron in the sample solution

s	= mass of flazasulfuron reference standard in the calibration solution (mg)

w   = mass of sample taken (mg)

P	= purity of flazasulfuron reference standard (g/kg)

FLAZASULFURON WATER DISPERSIBLE GRANULES

- 595/WG/M/-

1 Sampling. Take at least 100 g.

2 Identity test

2.1 HPLC. As for flazasulfuron technical 595/TC/M/2.1

#### 3 Flazasulfuron

#### As for flazasulfuron technical 595/TC/M/3 except

- (c) Sample preparation. Prepare sample solutions in duplicate for each sample. Grind the sample in a mortar in order to obtain a fine powder. Weigh (to the nearest 0.1 mg) sufficient sample (w mg) to contain about 40 mg flazasulfuron into a volumetric flask (100 ml). Add acetonitrile (about 75 ml) and sonicate for at least 15 minutes until complete dissolution. Allow the solutions to cool to ambient temperature and fill to the mark with acetonitrile (sample solutions S1 and S2). Mix thoroughly. Filter an aliquot of each prepared solution portion of each prepared solution through a 0.45 µm PTFE filter prior to analysis.

###### Typical Chromatograms 595/WG/M/3

Calibration solution of FLAZASULFURON

Sample solution of FLAZASULFURON 25 WG

Solution of solvent (acetonitrile)

Solution of FLAZASULFURON 25 WG BLANK FORMULATION (without a.s.)