PIRIMIPHOS-METHYL

Collaborative Study

Full Scale Collaborative Study 
for the 
Determination of
Pirimiphos-methyl AI and formulations, by 
Capillary Gas Chromatography

Report to CIPAC

by

Syngenta Crop Protection AG

Dr. A. McIntyre / Susanne De Benedictis

Im Breitenloh 5

CH-4333 Münchwilen

Switzerland

June 2011

Content 	page

1.	Participants	3

2.	Active Ingredient: General Information	5

3.	Samples	6

4.	Method	6

4.1	Scope	6

4.2	Principle	6

4.3	Procedure	6

5.	Remarks of the Participants	6

6.	Evaluation and Discussion	8

6.1	Data Review	8

6.2	Statistical Evaluation	8

7.	Conclusions	20

# Participants

In November 2010, Information Sheet No. 287 was sent out by the CIPAC Secretary inviting members to participate in a collaborative study on the determination of Pirimiphos-methyl, by capillary gas chromatography.

By the Middle of May 2011, 17 of the 26 respondents sent back their results.

Participants are listed in alphabetical order whereas lab numbers in the result tables were assigned, chronologically, based upon receipt of results.

| Benke, Lajos | Agricultural Office Directorate of Plant Protection and Soil Conservation 2481 Velence Hungary |
| --- | --- |
| Bernes, Albert Ing. | Walloon Agricultural Research Center 5030 Gembloux Belgium |
| Brennan, David Dr. | DuPont Crop Protection Newark, Delaware 19711-3507 United States |
| Christensen, Helena S Dr. | Cheminova A/S 7673 Harbooere Denmark |
| Çolak, İlknur Dr. | Ministry of Agriculture and Rural Affairs Directorate of Istanbul Provincial Agriculture Pesticide Quality Control Laboratuary Erenköy / İstanbul Turkey |
| De Mata, Maria Lourdes D. | OIC, Laboratory Services Devision National Pesticide Analytical Laboratory Bureau of Plant Industry Quezon City 1100 Philippines |
| Garvey, Jim Dr | Pesticide Control Laboratory Backweston Celbridge Co. Kildare Ireland |
| Grecu, Cornel | ALCHIMEX SA 010116 Bucharest-sector 1 Romania |
| Haustein, Michael Dr. | Currenta GmbH&Co. CUR-ANT Inprozesskontrolle 41538 Dormagen Germany |
| Ho, Ming-Hsun | TACTRI, Council of Agriculture Wufong, Taichung 41358 Taiwan (R.O.C) |
| Hosmani, Rajendra | Syngenta Bioscience Pvt. Ltd. Santa Monica Works Corlim, Ilhas, GOA 403110 India |
| Krste, Tashev | State Phytosanitary Laboratory 1000 Skopje, Republic of Macedonia |
| Lubyte, Jadvyga | Lithuanian Research Centre for Agriculture and Forestry Agrochemical Research Laboratory 50127 Kaunas Lithuania |
| Manso, Luis | Ministerio de Medio Ambiente, y Medio Rural y Marino Casiopea, Laboratorio Arbitral Agroalimentario 28023 Madrid Spain |
| Mojasevic, Milica | University of Belgrade Faculty of Agriculture, Pesticide Department 11081 Zemun Serbia |
| Nováková, Olga Dipl-Ing | State Phytosanitary Administration Pesticides Testing Laboratory 613 00 Brno Czech Republic |
| Olsen, Nongpanga | Office of Agricultural Research and Development Region1 A. Sansai Chiang Mai 50290 Thailand |

# Active Ingredient: General Information

Chemical name:	O-2-diethylamino-6-methylpyrimidin-4-yl O, O-dimethyl phosphorothioate

ISO common name:	Pirimiphos-methyl

CAS-Nr.:  		29232-93-7

Structure:

Molecular mass: 	305.3

Empirical formula:	C11H20N3O3PS

Activity:  		Insecticide

# Samples

Five test samples and two analytical standards were sent to the participants:

- Pirimiphos-methyl tech., sample 1

- Pirimiphos-methyl tech., sample 2

- Pirimiphos-methyl CS300, sample 1

- Pirimiphos-methyl CS300, sample 2

- Pirimiphos-methyl EC500, sample 1

Pirimiphos-methyl analytical standard, 99.2 % purity, and internal standard 4,4’-dimethoxy-benzophenone.

# Method

## Scope

Pirimiphos-methyl active ingredient content of technical grade active ingredient and associated formulations.

## Principle

Pirimiphos-methyl is determined by capillary gas chromatography using split injection mode, flame ionization detection, and internal standardization.

## Procedure

Each sample was analyzed using four independent determinations. The samples were analyzed on two different days with duplicate injections of two weighings per sample. Test and reference solutions were prepared fresh on each day. In order to calculate the response factor the mean of the response factor (double injection) before and after the samples were used (bracketing). The sample content was calculated using the mean value of the duplicate injections.

# Remarks of the Participants

Several participants provided comments about the method performance and also made a note of any deviations from the method:

Laboratory 1	Column is 30m x 0.32 mm i.d.; Oven conditions changed due to column difference: first ramp 35°C/min, second ramp 60°C/min with hold time of 4.4 min; all other conditions as per procedure provided

Laboratory 2	column: CP SIL 13CB (14 % phenyl, 86 % dimethyl- polyphenylmethylsiloxan) phase was used with dimensions of 25m x 0.32 mm i.d.

Temperature program: 35°C/min (ramp 1)

Carrier gas: hydrogen

Laboratory 3	Carrier gas: nitrogen

High concentration of Pirimiphos-methyl and internal standard resulting in response of 500 mV can cause increased contamination of liner, chemical trap and capillary. Therefore we recommend adjusting of concentrations to the response 100 – 200 mv.

Laboratory 4	column length: 30 m

Laboratory 5	carrier gas: nitrogen
flow rate. 1.0 ml/min

Laboratory 6	new column was used

Laboratory 7	no comments

Laboratory 8	no comments

Laboratory 9	column length: 30 m

Laboratory 10	column length: 30 m; but same chromatographic conditions

Laboratory 11	column length: 30 m;

carrier gas: nitrogen
split ratio 45:1 (due to total flow max 200 ml/min)

Laboratory 12	column length: 30 m

Laboratory 13	column length: 25 m; 
split ratio 50:1

Laboratory 14	column length: 25 m; 
carrier gas: nitrogen

Laboratory 15	column: DB-5 (5% Polyphenylmethylsiloxan) phase was used with dimensions of 30 m x 0.32 mm i.d.

Laboratory 16	no comments

Laboratory 17	column length: 30 m

Temperature program: first ramp 12.5°C/min, second ramp 20°C/min

# Evaluation and Discussion

## Data Review

The data obtained from each laboratory was reviewed to determine if there were any significant chromatography differences, from what was expected, which might affect the analytical results.

Visual examination of the chromatograms and data indicated no significant differences in all cases except Lab 9.

All method deviations, noted by the participants, were deemed not to affect the analytical results significantly.

Data review of Lab 9 showed a high variation of the area counts for both Pirimiphos-methyl and the internal standard. The area counts for the calibration solution varied randomly by a factor of approximately 20 throughout the whole sequence. This implied that the GC used was not operating satisfactorily. It was decided to exclude all data obtained from this laboratory for the statistical evaluation. The Lab was contacted to discuss these findings.

## Statistical Evaluation

The statistical evaluation of the data was accomplished following the “Guidelines for CIPAC Collaborative Study Procedures for Assessment of Performance of Analytical Methods”, according to DIN ISO 5725. Results reported by the laboratories and the statistical evaluation of these are listed in tables 1-3 and displayed in figures 1-5. These initial results are reported without any exclusion of outliers and/or stragglers.

The statistical evaluation in Table 3 shows that, without elimination of any outliers or stragglers, the between lab experimental Relative Reproducibility Standard Deviation (% RSDR) is below the calculated acceptable Horwitz value (% RSDR (Hor)) for the Pirimiphos-methyl technical samples 1 and 2, and for the EC500. For the CS300 samples, sample 1 has a result equal to the Horwitz criterion with sample 2 failing.

The data was examined for outliers and stragglers using Cochran’s test (within-lab variance), followed by Grubb’s test on the lab means (between lab variance). The tests were performed at an alpha level of 0.01 for outlier, and 0.05 for straggler detection.  Based on this procedure, the Cochran variance homogeneity test identified Laboratory 4 as one outlier for the EC500 data. The data of Laboratory 13, CS300-sample 1, was close to being designated a Cochran straggler with the same data indicating an outlier using Mandel’s k-statistic. The Mandel’s k-statistic was employed as a confirmation technique as this sample was on the borderline for the Cochran test. The Grubb’s test identified Laboratory 16 as an outlier, CS300 sample 2, and Laboratory 14 as a straggler for Pirimiphos-methyl technical sample 2. No iterations were required during the outlier/straggler detection process.

After elimination of any outliers, in accordance with both Cochran and Grubb’s test statistics, the evaluation in Table 4 showed that the between lab % RSDR is now below the acceptable % RSDR (Hor) value for the Pirimiphos-methyl CS300 sample 2. The data for the EC500 sample has been further improved although it was acceptable before the outlier was eliminated.

By excluding the data of Lab13 for the CS300 sample 1, which was close to being designated as a Cochran straggler, and subsequently reported as an outlier according to Mandel’s k-statistic, the between lab % RSDR falls below the acceptable % RSDR (Hor) value. This is illustrated within Table 5 where the statistical evaluation is reported after elimination of the Cochran and Grubbs outliers and this potential Cochran straggler. The results, post this final exclusion process, are displayed in figures 6-10.

#### Determination of Pirimiphos-methyl – no elimination of any outliers / stragglers

All results tabulated in table 1 and 2 are given in g/kg

Table 1 Results

|  | Pirimiphos-methyl tech. - sample 1 |  | Pirimiphos-methyl tech. -sample 2 |  | Pirimiphos-methyl CS300 -sample 1 |  | Pirimiphos-methyl CS300 -sample 2 |  | Pirimiphos-methyl EC500 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Day1 | Day2 | Day1 | Day2 | Day1 | Day2 | Day1 | Day2 | Day1 | Day2 |
| Lab 1 | 962.0 | 958.6 | 906.8 | 905.1 | 284.9 | 285.2 | 286.4 | 285.1 | 487.8 | 488.2 |
| Lab 2 | 963.4 | 956.5 | 910.8 | 903.3 | 282.0 | 282.7 | 281.3 | 283.7 | 487.6 | 486.1 |
| Lab 3 | 961.0 | 962.6 | 911.3 | 900.9 | 287.6 | 286.9 | 285.3 | 286.9 | 493.4 | 487.3 |
| Lab 4 | 959.7 | 959.3 | 904.1 | 908.3 | 275.3 | 284.0 | 272.8 | 283.8 | 516.5** | 490.7** |
| Lab 5 | 952.8 | 945.9 | 911.7 | 907.1 | 287.2 | 285.4 | 284.9 | 287.5 | 488.9 | 482.8 |
| Lab 6 | 958.1 | 955.8 | 906.8 | 904.8 | 292.0 | 300.6 | 292.1 | 300.2 | 496.1 | 502.2 |
| Lab 7 | 948.9 | 954.1 | 903.0 | 911.4 | 283.2 | 283.0 | 282.3 | 285.5 | 482.0 | 483.7 |
| Lab 8 | 956.3 | 952.1 | 906.8 | 902.0 | 287.8 | 283.9 | 284.3 | 281.6 | 489.2 | 487.8 |
| Lab 10 | 954.0 | 951.0 | 894.5 | 895.5 | 280.0 | 271.6 | 276.8 | 271.9 | 482.0 | 478.7 |
| Lab 11 | 954.7 | 957.3 | 900.5 | 904.3 | 288.9 | 283.4 | 289.8 | 286.9 | 486.9 | 487.3 |
| Lab 12 | 964.8 | 967.6 | 906.5 | 910.6 | 284.7 | 286.0 | 285.5 | 287.5 | 486.0 | 489.5 |
| Lab 13 | 960.6 | 959.8 | 904.8 | 906.3 | 267.4* | 283.5* | 274.6 | 282.2 | 490.2 | 488.5 |
| Lab 14 | 966.2 | 958.3 | 932.9 | 913.9 | 288.1 | 284.7 | 293.5 | 286.9 | 502.1 | 496.0 |
| Lab 15 | 967.7 | 966.3 | 915.1 | 916.0 | 277.3 | 280.3 | 272.6 | 275.4 | 487.4 | 485.9 |
| Lab 16 | 958.3 | 956.5 | 911.1 | 901.3 | 272.1 | 266.3 | 239.4 | 250.7 | 483.2 | 484.9 |
| Lab 17 | 956.2 | 963.3 | 894.6 | 899.9 | 283.2 | 281.9 | 282.3 | 285.1 | 493.3 | 496.6 |

*potential Cochran straggler (reported as outlier according to Mandel’s k-statistic)

** Cochran outlier

Table 2 Mean values

|  | Pirimiphos-methyl tech. - sample 1 | Pirimiphos-methyl tech. - sample 2 | Pirimiphos-methyl CS300 - sample 1 | Pirimiphos-methyl CS300 - sample 2 | Pirimiphos-methyl EC50 0 |
| --- | --- | --- | --- | --- | --- |
| Lab 1 | 960.3 | 906.0 | 285.1 | 285.8 | 488.0 |
| Lab 2 | 960.0 | 907.1 | 282.4 | 282.5 | 486.9 |
| Lab 3 | 961.8 | 906.1 | 287.3 | 286.1 | 490.4 |
| Lab 4 | 959.5 | 906.2 | 279.7 | 278.3 | 503.6 |
| Lab 5 | 949.4 | 909.4 | 286.3 | 286.2 | 485.9 |
| Lab 6 | 957.0 | 905.8 | 296.3 | 296.2 | 499.2 |
| Lab 7 | 951.5 | 907.2 | 283.1 | 283.9 | 482.9 |
| Lab 8 | 954.2 | 904.4 | 285.9 | 283.0 | 488.5 |
| Lab 10 | 952.5 | 895.0 | 275.8 | 274.4 | 480.4 |
| Lab 11 | 956.0 | 902.4 | 286.2 | 288.4 | 487.1 |
| Lab 12 | 966.2 | 908.6 | 285.4 | 286.5 | 487.8 |
| Lab 13 | 960.2 | 905.6 | 275.5 | 278.4 | 489.4 |
| Lab 14 | 962.3 | 923.4+ | 286.4 | 290.2 | 499.1 |
| Lab 15 | 967.0 | 915.6 | 278.8 | 274.0 | 486.7 |
| Lab 16 | 957.4 | 906.2 | 269.2 | 245.1++ | 484.1 |
| Lab 17 | 959.8 | 897.3 | 282.6 | 283.7 | 495.0 |

+ Grubbs straggler

++ Grubbs outlier

Table 3 Summary of the statistical evaluation

No elimination of any outliers / stragglers

|  | Pirimiphos-methyl tech. - sample 1 | Pirimiphos-methyl tech. - sample 2 | Pirimiphos-methyl CS300 - sample 1 | Pirimiphos-methyl CS300 - sample 2 | Pirimiphos-methyl EC500 |
| --- | --- | --- | --- | --- | --- |
| xm | 958.4 | 906.6 | 282.8 | 281.4 | 489.7 |
| L | 16 | 16 | 16 | 16 | 16 |
| sr | 3.07 | 5.07 | 4.28 | 3.96 | 5.20 |
| sL | 4.42 | 5.35 | 5.39 | 10.88 | 5.22 |
| sR | 5.39 | 7.37 | 6.89 | 11.58 | 7.37 |
| r | 8.60 | 14.21 | 11.99 | 11.09 | 14.55 |
| R | 15.08 | 20.64 | 19.28 | 32.42 | 20.63 |
| RSDr | 0.32 | 0.56 | 1.51 | 1.41 | 1.06 |
| RSDR | 0.56 | 0.81 | 2.43* | 4.11** | 1.50 |
| RSDR(Hor) | 2.01 | 2.03 | 2.42 | 2.42 | 2.23 |

*Borderline with Horwitz limit

**fails Horwitz limit

xm		=	overall sample mean

L		=	number of laboratories

sr 		=	repeatability standard deviation

RSDr		=	relative repeatability standard deviation

r		=	repeatability limit

sR		=	reproducibility standard deviation

RSDR		=	relative reproducibility standard deviation

R		=	reproducibility limit

sL		=	“pure” between laboratory standard deviation

RSDR(Hor)	=	Horwitz relative reproducibility standard deviation

Fig. 1

Pirimiphos-methyl tech. – sample 1

R limits

r limits

Mean

Fig. 2

Pirimiphos-methyl tech. – sample 2

R limits

r limits

Mean

Fig. 3

Pirimiphos-methyl CS300 – sample 1

R limits

r limits

Mean

Fig. 4

Pirimiphos-methyl CS300 – sample 2

R limits

r limits

Mean

Fig. 5

Pirimiphos-methyl EC500 – sample 1

R limits

r limits

Mean

Table 4 Summary of the statistical evaluation

Post elimination of outliers according to Cochran and Grubb’s statistics

|  | Pirimiphos-methyl tech. - sample 1 | Pirimiphos-methyl tech. - sample 2 | Pirimiphos-methyl CS300 - sample 1 | Pirimiphos-methyl CS300 - sample 2 | Pirimiphos-methyl EC500 |
| --- | --- | --- | --- | --- | --- |
| xm | 958.4 | 906.6 | 282.8 | 283.8 | 488.7 |
| L | 16 | 16 | 16 | 15 | 15 |
| sr | 3.07 | 5.07 | 4.28 | 3.53 | 2.57 |
| sL | 4.42 | 5.35 | 5.39 | 5.32 | 5.05 |
| sR | 5.39 | 7.37 | 6.89 | 6.39 | 5.67 |
| r | 8.60 | 14.21 | 11.99 | 9.89 | 7.20 |
| R | 15.08 | 20.64 | 19.28 | 17.88 | 15.88 |
| RSDr | 0.32 | 0.56 | 1.51 | 1.24 | 0.53 |
| RSDR | 0.56 | 0.81 | 2.43 | 2.25* | 1.16** |
| RSDR(Hor) | 2.01 | 2.03 | 2.42 | 2.42 | 2.23 |

*  Elimination of Grubbs outlier brings the value below the Horwitz acceptance criterion

** Elimination of Cochran outlier improves the data further although value was acceptable before elimination

xm		=	overall sample mean

L		=	number of laboratories

sr 		=	repeatability standard deviation

RSDr		=	relative repeatability standard deviation

r		=	repeatability limit

sR		=	reproducibility standard deviation

RSDR		=	relative reproducibility standard deviation

R		=	reproducibility limit

sL		=	“pure” between laboratory standard deviation

RSDR(Hor)	=	Horwitz relative reproducibility standard deviation

Table 5 Summary of the statistical evaluation

Post elimination of Cochran and Grubb’s outliers and the potential Cochran straggler which was identified as an outlier according to Mandel’s k statistic

|  | Pirimiphos-methyl tech. - sample 1 | Pirimiphos-methyl tech. - sample 2 | Pirimiphos-methyl CS300 - sample 1 | Pirimiphos-methyl CS300 - sample 2 | Pirimiphos-methyl EC500 |
| --- | --- | --- | --- | --- | --- |
| xm | 958.4 | 906.6 | 283.3 | 283.8 | 488.7 |
| L | 16 | 16 | 15 | 15 | 15 |
| sr | 3.07 | 5.07 | 3.30 | 3.53 | 2.57 |
| sL | 4.42 | 5.35 | 5.60 | 5.32 | 5.05 |
| sR | 5.39 | 7.37 | 6.50 | 6.39 | 5.67 |
| r | 8.60 | 14.21 | 9.25 | 9.89 | 7.20 |
| R | 15.08 | 20.64 | 18.20 | 17.88 | 15.88 |
| RSDr | 0.32 | 0.56 | 1.17 | 1.24 | 0.53 |
| RSDR | 0.56 | 0.81 | 2.29* | 2.25 | 1.16 |
| RSDR(Hor) | 2.01 | 2.03 | 2.42 | 2.42 | 2.23 |

* after elimination of the potential Cochran straggler (Mandel’s k-statistic outlier)  the Horwitz criterion is fulfilled.

xm		=	overall sample mean

L		=	number of laboratories

sr 		=	repeatability standard deviation

RSDr		=	relative repeatability standard deviation

r		=	repeatability limit

sR		=	reproducibility standard deviation

RSDR		=	relative reproducibility standard deviation

R		=	reproducibility limit

sL		=	“pure” between laboratory standard deviation

RSDR(Hor)	=	Horwitz relative reproducibility standard deviation

Fig. 6

Pirimiphos-methyl tech. – sample 1

R limits

r limits

Mean

Fig. 7

Pirimiphos-methyl tech. – sample 2

R limits

r limits

Mean

Fig. 8

Pirimiphos-methyl CS300 – sample 1

R limits

r limits

Mean

Fig. 9

Pirimiphos-methyl CS300 – sample 2

R limits

r limits

Mean

Fig. 10

Pirimiphos-methyl EC500 – sample 1

R limits

r limits

Mean

# Conclusions

A total of 17 different laboratories have participated in this full scale collaborative study.

Upon review of the chromatographic data all results appear to be valid with the exception of laboratory 9 where a high variation of area counts, for repeat injections, indicates that the instrument used was not operating satisfactorily. Hence the results of this lab could not be considered as valid data and have been excluded from the calculations. The laboratory was informed and discussions held to identify potential causes for the problem.

Without elimination of any outliers or stragglers the between lab experimental Relative Reproducibility Standard Deviation ( % RSDR ) is below the calculated acceptable value based on the Horwitz curve calculation (% RSDR (Hor)) for the Pirimiphos-methyl technical samples 1 and 2 as well as for the EC500 sample. For the CS300 samples, sample 1 is reported equal to the Horwitz criterion with sample 2 failing (Table 3).

After elimination of the identified Cochran and Grubb’s outliers the % RSDR for CS300 sample 2 passes the Horwitz criterion, and the value for the EC500 sample has been improved although it was deemed acceptable without the elimination of the outlier (Table 4).

After elimination of the possible Cochran straggler, which is also reported as a Mandel’s k–statistic outlier within the CS300 sample 1 data set, the % RSDR falls below the acceptable % RSDR (Hor) value (Table 5).

The fact that both the Technical Grade Active Ingredient (TGAI) samples and the EC formulation sample passed the Horwitz criterion, without removal of any outliers/stragglers, and the CS samples returning a borderline pass and failure for samples 1 and 2 respectively, provides a clear indication that the sample preparation and homogeneity is a key element. The core chromatographic method is common to all samples. The importance of well homogenized samples prior to sample preparation was also seen during a DAPA (Deutscher Arbeitskreis für Pflanzenschutz-Analytik) small scale collaborative trial and therefore the sample preparation section in the method was emphasised within the Draft CIPAC method.

The overall indication is that the higher between-laboratory experimental Relative Reproducibility Standard Deviation (% RSDR) for the CS300 formulation is a result of poor sample homogeneity rather than the experimental conditions of the method.

Using larger sample amounts was considered as a potential way of minimising the homogeneity issue but this would be a balance as excessive amount of solvent and an extra dilution step may have to be incorporated. Increased HSE issues, due to increased solvent volumes, and the analytical error introduced by an extra dilution step were not seen as improvements upon a method which we believe is operating satisfactorily when the sample homogeneity is addressed.

In order to improve upon this homogeneity issue Syngenta would recommend the inclusion of a more specific mixing instruction than is currently within the method.

The current instruction is as follows:-

‘Homogenize the test sample carefully’

The proposal would be to include a phrase similar to the following:-

‘Thoroughly homogenize the formulation, as received in the original 1L sales pack, by inverting/shaking the bottle for 2 minutes or by using a mechanical stirrer (Syngenta used a Turbula 3D-rotary mixer for 10 minutes) Do not decant a smaller sample or remove the test sample without thorough homogenization.’

Taking into account the number of participating Laboratories, 16*, and the fact that the samples with a more homogeneous nature, TGAI and EC500, passed the acceptance criterion easily without exclusion of any outliers/stragglers, and the less homogeneous CS300 samples passed the acceptance criterion after elimination of only 1 outlier/straggler per sample, we consider this method to be suitable and recommend accepting it as a provisional CIPAC method for the determination of Pirimiphos-methyl in TGAI and formulations EC and CS .

*Laboratory 9 was discounted at the beginning due to gross discrepancies being visible within the raw data