SYRICI

Pyraoxystrobin 964

CIPAC Collaborative Trial

CIPAC 4937/R, full scale study

CIPAC Collaborative Study of a

High Performance Liquid Chromatographic Analysis of Pyraoxystrobin Technical Material and Formulated Products

By

Professor Zhao Xinxin and

Ms Wang Haixia

Shenyang Research Institute of Chemical Industry Co., Ltd

Pesticides Test Laboratory

No.8, Shenliao Dong Road, Tiexi District, Shenyang City,

P. R. of China

Presented at the CIPAC meeting in Liège, Belgium

June, 2014

Pages: 30

Table of Contents

1.	List of Participants	3

2.	General Information	4

3.	Distribution of Samples	5

4.	Procedure	5

4.1.	Outline of Method	5

4.2.	Program of Work	6

5.	Analytical Methods	6

5.1.	Analytical Conditions	6

5.2.	Deviations from the Analytical Method	8

5.3.	Remarks about the Analytical Method	10

6.	Statistical Evaluation	11

7.	Results	12

8.	Summary of the results	27

9.	Statistical formulas	28

10.	Discussion	29

# List of Participants

26 laboratories took part in the collaborative study:

| NAME | ORGANIZATION | COUNTRY |
| --- | --- | --- |
| Banu Ioana Silvia | Central Phytosanitary Laboratory | Romania |
| Helen Karasali | Benaki Phytopathological Institute | Greece |
| Peter Derevyanko Iwan Orgei | Frandesa Co. Ltd. | Belarus |
| Liliya Petrunova | Central Laboratory for Chemical Testing and Control, Bulgarian Food Safety Agency | Bulgaria |
| Cornel Grecu | Alchimex SA | Romania |
| Eva Jacobsen | Danish Technological Institute, Life Science, Laboratory for Chemistry and Microbiology | Denmark |
| Dr. Alice Egger | Currenta GmbH & Co. OHG | Germany |
| Lajos Benke | National Food Chain Safety Office, Pesticide Analytical Laboratory Velence | Hungary |
| Ing. T.C. de Rijk | RIKILT-DLO | Netherlands |
| Brenda Checa Orrego / José Manuel Bocanegra | Ministerio de Desarrollo Agropecuario, Dirección Nacional de Sanidad Vegetal, Laboratorio de Control de Calidad de Plaguicidas y Fertilizantes | Panama |
| Ms Susan Marais | Pesticide Analytical Technology | South Africa |
| Angela Santilio | National Institute of Health - Department of Enviromental and Primary Prevention - Pesticide Unit | Italy |
| Elizabeth Carranza de Aguila / Billy Franz Arteaga | Laboratorio Control de Calidad de Plaguicidas MAG-OIRSA | El Salvador |
| SELMA BELART | Laboratorio Plaguicidas - SENASA | Argentina |
| Hrybova Nataliia | Ukrainian Laboratory of Quality and Safety of Agricultural Products (ULQSAP) of the National University of Life and Environmental Sciences of Ukraine (NUBiP of Ukraine) | Ukraine |
| Milan Damjanovic | Chemical Agrosava | Serbia |
| Krste Tashev | State Phytosanitary Laboratory (Ministry of Agriculture, Forestry and Water Economy) | Republic of Macedonia |
| Sung Lim, Ko | FACT (Foundation of Agri. Tech, Commercialization & Transfer) | Republic of Korea |
| Tiechun Chen | Quality Control Division, Institute for the Control of Agrochemicals, Ministry of Agriculture | China |
| Lucija Janeš | Agricultural Institute OF Slovenia (Kmetijski inštitut Slovenije) | Slovenia |
| Gareth Mann | Absolute Science | South Africa |
| Yusuf Vohra, Dr. S.Y. Pandey and Dr. Nadeem Ahmad Khan | Jai Research Foundation | India |
| Trevor Bowen | Bayer CropScience AG | Germany |
| Rose Feng | Jiangsu Rotam Chemistry Co., Ltd. | China |
| Dr. A.Ramesh | International Institute of Biotechnology and Toxicology | India |
| Yue Wang | Nutrichem Laboratory Co., Ltd. | China |

# General Information

Pyraoxystrobin

964/TC/M/

ISO Common name    Pyraoxystrobin

Chemical name        methyl (2E)-2-(2-{[3-(4-chlorophenyl)-1-

methylpyrazol-5-yl]oxymethyl}phenyl)-3-

methoxyacrylate (IUPAC)

methyl (αE)-2-[[[3-(4-chlorophenyl)-1-methyl-1H-

pyrazol-5-yl]oxy]methyl]-α-(methoxymethylene)

benzeneacetate (CA; 862588-11-2)

Empirical formula      C22H21ClN2O4

RMM                 412.87

m.p.                  129.6℃

v.p.                   1.22×10-8 Pa at 20℃

Solubility             In water0.03 mg/L at 20℃; High solubility in

N,N-Dimethylformamide,acetone, ethyl acetate,

toluene, acetonitrile

Description            The pure material is a white, odourless solid

Stability               Not stable in alcohol to heat

Formulations           Suspension concentrates (SC)

# Distribution of Samples

The following samples were provided to the partcipants:

Reference standard of Pyraoxystrobin                    1.0 g

Lot No. 20130130, purity: 997 g/kg

Pyraoxystrobin technical (TC1)                          3.0 g

Batch No. 20130205001, approx. 970 g/kg

Pyraoxystrobin technical (TC2)                          3.0 g

Batch No. 20130211001, approx. 971 g/kg

Pyraoxystrobin 20% suspension concentrate (SC1)        60 mL

Batch No. 20130226001, approx. 202 g/kg

Pyraoxystrobin 20% suspension concentrate (SC2)        60 mL

Batch No. 20130308001, approx. 203 g/kg

Pyraoxystrobin 20% suspension concentrate (SC3)        60 mL

Batch No. 20130315001, approx. 208 g/kg

# Procedure

## Outline of Method

Pyraoxystrobin in the test substance is determined by HPLC method with a C18-reverse phase column and UV-detection at 280 nm. The analyte solution contains about 40 mg of pyraoxystrobin in 100 mL solution.

## Program of Work

We requested the collaborators to:

1) conduct duplicate determinations on two different days for each of the five samples;

2) inject each sample solution in duplicate and calculate the mean value;

3) check equilibration of the system before the determination;

4) describe operating conditions in detail; and

5) attach the typical chromatograms for the five samples.

# Analytical Methods

## Analytical Conditions

| Lab | Liquid chromatograph integrator | Column | Mobile phase | Flow rate (mL/min) | Column temp(℃) |
| --- | --- | --- | --- | --- | --- |
| Proposed Conditions |  | Agilent Zorbax SB-C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 1 | Varian Pro Star, Model 240 | Prodigy ODS, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 2 | UFLC Shimadzu LC-20AD | LIchroCart Purospher RP-C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.5 | 30 |
| 3 | Agilent 1260 Infinity | Agilent Zorbax SB-C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 4 | LaChrom Elite VWR-Hitachi | Agilent Zorbax ODS, 5µm 250×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 5 | Waters | Thermo Hypersil ODS, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.0 | 30 |
| 6 | Agilent 1260 Infinity | Phenomenex Kinetex C18, 5µm 150×4.6mm (i.d.) | A: acetonitrile B:water Time A% B% 0 40 60 4.8 40 60 4.9 20 80 7.0 20 80 7.2 40 60 10.0 40 60 | 1.3 | 30 |
| 7 | Shimadzu LC 2010 | Agilent Zorbax SB-C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 8 | Dionex HPLC system | Agilent Zorbax SB-C18, 5µm 250×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.0 | 25 |
| 9 | Jasco PU-2089 plus | ZORBAX Eclipse XDB-C18, 3.5µm, 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | / |
| 10 | Agilent 1260 HPLC Technologies | Agilent Eclipse XDB-C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 11 | Agilent 1200 HPLC | Zorbax Eclipse XDB-C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | ambinet |
| 12 | Perkin Elmer series 200 | Agilent Zorbax SB-C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 13 | Agilent 1200 series | Agilent Zorbax SB-C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 14 | Shimadzu LC 20A | Phenomenex Luna C18, 5µm 250×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 15 | Dionex -Ultimate 3000 | Acclaim C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 16 | HPLC Agilent 1220 Infinity LC | Agilent Zorbax SB-C8, 3.5µm 150×4.6mm (i.d.) | acetonitrile – water, 600 + 400 (v/v) | 1.3 | Ambient |
| 17 | Varian Pro Star | Perkin Elmer C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 18 | Shimadzu LC-20A | Unison UK-C18, 3µm 100×2.0mm (i.d.) | Acetonitrile-water, 700+400(v/v) | 0.3 | 40 |
| 19 | Agilent 1200 HPLC | Agilent Zorbax SB-C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 20 | HPLC Agilent Technologies 1200 Series | BIA d.o.o., Lichrospher 100 RP-C18, 5µm 250×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 21 | Agilent Infinity 1260 | Agilent Eclipse Plus 3.5µm, 100×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 22 | Shimadzu LC-2010 AHT | Agilent Eclipse XDB, 3.5µm, 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.0 | 30 |
| 23 | Agilent series 1260 Infinity | Phenomenex Luna C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 24 | Shimadzu LC20A | Agilent Eclipse XDB-C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 25 | Agilent-1200 Series | Kromasil-100 C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |
| 26 | Agilent 1260 | Agilent Zorbax SB-C18, 5µm 150×4.6mm (i.d.) | acetonitrile - water, 600+400 (v/v) | 1.3 | 30 |

## Deviations from the Analytical Method

Lab 1:   Column: Prodigy ODS, 5µm, 150×4.6mm (i.d.).

Lab 2:   Column: LIchroCart Purospher RP-C18, 5µm, 150×4.6mm (i.d.);

Flow rate: 1.5 mL/min.

Lab 3:   No deviations.

Lab 4:   Column: Agilent Zorbax ODS, 5µm, 250×4.6mm (i.d.).

Lab 5:   Column: Thermo Hypersil ODS, 5µm, 150×4.6mm (i.d.);

Flow rate: 1.0 mL/min.

Lab 6:   Column: Phenomenex Kinetex C18, 5µm, 150×4.6mm (i.d.);

Injection volume: 10µL;

Mobile phase: 60% acetonitrile and 40% water (0-4.8 min), 80%

acetonitrile and 20% water (4.9-7.0 min), 60% acetonitrile and

40% water (7.2-10.0 min).

Lab 7:   No deviations.

Lab 8:   Column: Agilent Zorbax SB-C18, 5µm, 250×4.6mm (i.d.);

Column temperature: 25℃; Flow rate: 1.0 mL/min.

Lab 9:   Column: Zorbax Eclipse XDB-C18, 3.5µm, 150×4.6mm (i.d.).

Lab 10:  Column: Agilent Eclipse XDB-C18, 5µm, 150×4.6mm (i.d.).

Lab 11:  Column: Zorbax Eclipse XDB-C18, 5µm, 150×4.6mm (i.d.);

Column temperature: ambient;

Reduced sonication time to 2 minutes;

Weigh approximately 20mg in 50ml volumetric flasks instead of

40mg in 100ml flasks.

Lab 12:  No deviations.

Lab 13:  No deviations.

Lab 14:  Column: Phenomenex Luna C18, 5µm, 150×4.6mm (i.d.).

Lab 15:  Column: Acclaim C18, 5µm, 150×4.6mm (i.d.).

Lab 16:  Column: Agilent Zorbax SB-C8, 3.5µm, 150×4.6mm (i.d.);

Column temperature: ambient;

Lab 17:  Column: Perkin Elmer C18, 5µm, 150×4.6mm (i.d.).

Lab 18:  Column: Unison UK-C18, 3µm, 100×2.0mm (i.d.);

Mobile phase: acetonitrile – water, 700+400 (v/v);

Column temperature: 40℃; Flow rate: 0.3 mL/min.

Lab 19:  No deviations.

Lab 20:  Column: Lichrospher 100 RP-C18, 5µm, 250×4.6mm (i.d.).

Lab 21:  Column: Agilent Eclipse Plus, 3.5µm, 100×4.6mm (i.d.).

Lab 22:  Column: Agilent Eclipse XDB, 3.5µm, 150×4.6mm (i.d.);

Flow rate: 1.0 mL/min.

Lab 23:  Column: Phenomenex Luna C18, 5µm, 150×4.6mm (i.d.).

Lab 24:  Column: Agilent Eclipse XDB-C18, 5µm, 150×4.6mm (i.d.).

Lab 25:  Column: Kromasil-100C18, 5µm, 150×4.6mm (i.d.).

Lab 26:  No deviations.

## Remarks about the Analytical Method

Lab 1:   Is a stable and robust method.

Lab 2:   The method is perfect. There was no problem during analysis.

Lab 3:   Recommend to choose a wavelength equal to 254 nm,

unnecessarily in this case, a maximum sensitivity of the detector.

Thanks to this, you can reduce the injection volume, thus

reducing the possible influence of the matrix in formulations.

It is possible to use approximately 5 minutes instead of the

recommended 15 minutes to dissolve the sample in an ultrasonic

bath. It is enough to completely dissolve the weighed sample.

Lab 4:   No remarks.

Lab 5:   This method is ok but can be improved by using of internal

standard.

Lab 6:   No remarks.

Lab 7:   No remarks.

Lab 8:   No remarks.

Lab 9:   It would be recommended to emulsify the SC samples with 10%

water and then to dissolve in acetonitrile. It would be

recommended to quantify the samples using a calibration curve.

Lab 10:  No remarks.

Lab 11:  No remarks.

Lab 12:  No remarks.

Lab 13:  A similar column could be used with the same flow.

Lab 14:  It would be better to increase the temperature by 5 or 10°C.

Lab 15:  If necessary to confirm the identification picoxystrobin pesticide

formulation we could use HPLC / MS / MS.

Lab 16:  Retention time of Pyraoxystrobin was approximately 4.2 min.

Lab 17:  No remarks.

Lab 18:  No remarks.

Lab 19:  No remarks.

Lab 20:  No remarks.

Lab 21:  No remarks.

Lab 22:  Method is good and reproducible.

Lab 23:  No remarks.

Lab 24:  The method can be repeated.

Lab 25:   After sonication the sample solutions pertaining to SC

formulations were Filtered (as it was slightly turbid) through

0.45µm syringe filter before injection.

Lab 26:  No remarks.

# Statistical Evaluation

Samples were sent to 26 laboratories. All of them sent back results.

The statistical evaluation of the data was done following DIN ISO 5725 and “Guidelines for CIPAC Collaborative Study Procedure for Assessment of Performance of Analytical Methods”.

In Lab 4, the response factor for the calibration solution B differed much more than 1.0% from that for the calibration solution A, which did not comply with the procedure of HPLC method of pyraoxystrobin for CIPAC Full Scale Collaborative Trial. So the data of the Lab 4 was invalid and excluded from the statistical evaluation.

The assay results obtained by the collaborators and the statistical evaluation are reported in Tables 1 through 5, and in Figures 1-1 through 5-2. Formulas used are listed in sector 9, page 28.

# Results

Table 1: Pyraoxystrobin TC1 (Batch No: 20130205001)

| Lab | Day 1 |  | Day 2 |  | Mean g/kg | Std. Dev. |
| --- | --- | --- | --- | --- | --- | --- |
|  | A | B | A | B |  |  |
| 1 | 977.63 | 979.66 | 973.69 | 971.73 | 975.68 | 3.615 |
| 2 | 974.96 | 974.99 | 973.51 | 978.55 | 975.50 | 2.146 |
| 3 | 974.82 | 978.7 | 971.68 | 971.64 | 974.21 | 3.344 |
| 4 | 1071.36 | 1037.23 | 951.89 | 966.86 | 1006.84 | 56.875 |
| 5 | 966.44 | 972 | 982.79 | 972.06 | 973.32 | 6.840 |
| 6 | 974.4 | 979.55 | 966.86 | 976.31 | 974.28 | 5.384 |
| 7 | 976.91 | 973.82 | 973.03 | 978.7 | 975.62 | 2.652 |
| 8 | 970.87 | 971.46 | 970.74 | 971.97 | 971.26 | 0.568 |
| 9 | 969.52 | 969.67 | 968.5 | 967.05 | 968.69 | 1.208 |
| 10 | 962.16 | 977.42 | 970.71 | 969.6 | 969.97 | 6.250 |
| 11 | 977.58 | 979.95 | 974.59 | 975.94 | 977.02 | 2.307 |
| 12 | 965.65 | 961.09 | 965.99 | 987.06 | 969.95 | 11.625 |
| 13 | 963.43 | 976.52 | 936.28 | 979.76 | 964.00 | 19.781 |
| 14 | 973.47 | 966.03 | 966.53 | 958.35 | 966.10 | 6.180 |
| 15 | 981.81 | 981.47 | 983.57 | 981.6 | 982.11 | 0.982 |
| 16 | 1020.46 | 1023.6 | 993.96 | 1007.01 | 1011.26 | 13.592 |
| 17 | 947.11 | 967.2 | 980.73 | 981.11 | 969.04 | 15.986 |
| 18 | 976.04 | 964.08 | 957.85 | 956.52 | 963.62 | 8.910 |
| 19 | 978.07 | 976.03 | 981.84 | 973.75 | 977.42 | 3.433 |
| 20 | 986.01 | 986.2 | 985.67 | 983.95 | 985.46 | 1.029 |
| 21 | 985.96 | 959.26 | 976.74 | 977.68 | 974.91 | 11.226 |
| 22 | 950.97 | 954.1 | 958.74 | 959.09 | 955.73 | 3.901 |
| 23 | 969.61 | 960.9 | 969.49 | 959.76 | 964.94 | 5.344 |
| 24 | 964.33 | 971.19 | 973.1 | 976.17 | 971.20 | 5.017 |
| 25 | 970.28 | 970.64 | 971.68 | 970.49 | 970.77 | 0.623 |
| 26 | 975.65 | 970.99 | 971.48 | 978.64 | 974.19 | 3.629 |

Lab 4: Invalid data due to noncompliance with Pyraoxystrobin method.

Lab 13 and Lab 17: Outlier according to Cochran Test.

Lab 16: Outlier according to Grubbs Test.

Fig.1-1 : Pyraoxystrobin TC1 (Batch No: 20130205001)   All labs

Fig.1-2: Pyraoxystrobin TC1 (Batch No: 20130205001)   Labs 4, 13, 16 and 17 excluded

Table 2: Pyraoxystrobin TC2 (Batch No: 20130211001)

| Lab | Day 1 |  | Day 2 |  | Mean g/kg | Std. Dev. |
| --- | --- | --- | --- | --- | --- | --- |
|  | A | B | A | B |  |  |
| 1 | 978.3 | 979.73 | 971.04 | 972.39 | 975.37 | 4.290 |
| 2 | 966.84 | 969.65 | 963.52 | 972.92 | 968.23 | 4.005 |
| 3 | 975.9 | 974.88 | 967.07 | 969.34 | 971.80 | 4.271 |
| 4 | 1089.65 | 1057.25 | 1076.93 | 1087.31 | 1077.79 | 16.324 |
| 5 | 982.12 | 972.55 | 970.87 | 975.35 | 975.22 | 4.956 |
| 6 | 982.35 | 981.45 | 960.17 | 973.47 | 974.36 | 10.267 |
| 7 | 976.73 | 973.76 | 978.13 | 975.94 | 976.14 | 1.827 |
| 8 | 970.75 | 971.36 | 971.59 | 972.22 | 971.48 | 0.607 |
| 9 | 955.95 | 955.4 | 958.86 | 958.21 | 957.11 | 1.687 |
| 10 | 962.74 | 969.15 | 960.81 | 967.8 | 965.13 | 3.986 |
| 11 | 973.45 | 976.48 | 974.44 | 979.03 | 975.85 | 2.467 |
| 12 | 972.56 | 965.01 | 1028.39 | 997.13 | 990.77 | 28.583 |
| 13 | 977.98 | 972.54 | 974.03 | 965.57 | 972.53 | 5.177 |
| 14 | 965.89 | 975.83 | 968.59 | 966.93 | 969.31 | 4.487 |
| 15 | 984.99 | 985.87 | 986.1 | 989.24 | 986.55 | 1.856 |
| 16 | 1024.28 | 1013.89 | 1007.32 | 1000.88 | 1011.59 | 9.988 |
| 17 | 977.63 | 941.27 | 981.16 | 982.51 | 970.64 | 19.689 |
| 18 | 967.35 | 947.13 | 961.52 | 962.98 | 959.75 | 8.767 |
| 19 | 961.44 | 966.35 | 973.82 | 967.81 | 967.36 | 5.099 |
| 20 | 989.46 | 975.46 | 983.55 | 983.08 | 982.89 | 5.740 |
| 21 | 971.93 | 977.76 | 968.84 | 977.51 | 974.01 | 4.373 |
| 22 | 952.03 | 953.8 | 966.28 | 960.92 | 958.26 | 6.585 |
| 23 | 971.47 | 974.64 | 971.91 | 970.1 | 972.03 | 1.903 |
| 24 | 971.65 | 971.2 | 971.89 | 971.87 | 971.65 | 0.321 |
| 25 | 974.39 | 976.34 | 976.29 | 977.12 | 976.04 | 1.161 |
| 26 | 972.13 | 973.02 | 979.53 | 974.12 | 974.70 | 3.321 |

Lab 4: Invalid data due to noncompliance with Pyraoxystrobin method.

Lab 12 and Lab 17: Outlier according to Cochran Test.

Lab 16: Outlier according to Grubbs Test.

Fig.2-1: Pyraoxystrobin TC2 (Batch No: 20130211001)   All labs

Fig.2-2: Pyraoxystrobin TC2 (Batch No: 20130211001)   Labs 4, 12, 16 and 17 excluded

Table 3: Pyraoxystrobin SC1 (Batch No: 20130226001)

| Lab | Day 1 |  | Day 2 |  | Mean g/kg | Std. Dev. |
| --- | --- | --- | --- | --- | --- | --- |
|  | A | B | A | B |  |  |
| 1 | 204.22 | 206.09 | 204.44 | 204.38 | 204.78 | 0.877 |
| 2 | 206.40 | 202.22 | 205.73 | 201.02 | 203.84 | 2.627 |
| 3 | 203.44 | 205.07 | 203.63 | 204.41 | 204.14 | 0.750 |
| 4 | 219.89 | 207.63 | 200.51 | 208.43 | 209.12 | 8.017 |
| 5 | 211.73 | 209.77 | 214.64 | 211.95 | 212.02 | 2.001 |
| 6 | 198.08 | 198.16 | 198.18 | 198.19 | 198.15 | 0.050 |
| 7 | 203.21 | 202.33 | 202.04 | 203.67 | 202.81 | 0.758 |
| 8 | 202.88 | 202.34 | 202.20 | 201.27 | 202.17 | 0.669 |
| 9 | 188.33 | 187.83 | 189.14 | 188.03 | 188.33 | 0.576 |
| 10 | 196.40 | 195.31 | 202.18 | 202.40 | 199.07 | 3.743 |
| 11 | 201.35 | 201.60 | 200.70 | 200.25 | 200.98 | 0.614 |
| 12 | 201.63 | 200.90 | 202.34 | 203.41 | 202.07 | 1.069 |
| 13 | 203.47 | 204.64 | 202.72 | 206.26 | 204.27 | 1.543 |
| 14 | 195.75 | 195.86 | 198.97 | 198.28 | 197.22 | 1.653 |
| 15 | 213.03 | 216.92 | 215.26 | 219.29 | 216.13 | 2.644 |
| 16 | 220.00 | 222.04 | 217.23 | 217.17 | 219.11 | 2.358 |
| 17 | 210.26 | 210.53 | 206.48 | 204.99 | 208.07 | 2.761 |
| 18 | 200.26 | 198.73 | 201.29 | 200.66 | 200.24 | 1.089 |
| 19 | 204.68 | 204.00 | 201.14 | 200.79 | 202.65 | 1.973 |
| 20 | 201.23 | 202.20 | 202.03 | 203.27 | 202.18 | 0.839 |
| 21 | 202.36 | 201.48 | 199.15 | 200.43 | 200.86 | 1.384 |
| 22 | 198.48 | 195.94 | 198.85 | 203.48 | 199.19 | 3.140 |
| 23 | 203.60 | 205.46 | 190.15 | 200.03 | 199.81 | 6.823 |
| 24 | 200.98 | 200.87 | 200.27 | 200.45 | 200.64 | 0.337 |
| 25 | 200.41 | 201.09 | 201.00 | 200.23 | 200.68 | 0.427 |
| 26 | 201.06 | 200.98 | 201.12 | 200.12 | 200.82 | 0.470 |

Lab 4: Invalid data due to noncompliance with Pyraoxystrobin method.

Lab 9: Outlier according to Dixon Test.

Lab 16: Outlier due to deviation from the method (overlapping peaks).

Lab 23: Outlier according to Cochran Test.

Fig.3-1: Pyraoxystrobin SC1 (Batch No: 20130226001)   All labs

Fig.3-2: Pyraoxystrobin SC1 (Batch No: 20130226001)    Labs 4, 9, 16 and 23 excluded

Table 4: Pyraoxystrobin SC2 (Batch No: 20130308001)

| Lab | Day 1 |  | Day 2 |  | Mean g/kg | Std. Dev. |
| --- | --- | --- | --- | --- | --- | --- |
|  | A | B | A | B |  |  |
| 1 | 207.95 | 207.47 | 205.48 | 206.35 | 206.81 | 1.113 |
| 2 | 209.85 | 213.23 | 206.13 | 212.71 | 210.48 | 3.259 |
| 3 | 204.17 | 204.06 | 204.12 | 205.13 | 204.37 | 0.509 |
| 4 | 210.54 | 216.35 | 201.35 | 220.42 | 212.17 | 8.272 |
| 5 | 201.01 | 200.07 | 206.06 | 205.89 | 203.26 | 3.162 |
| 6 | 197.83 | 197.56 | 198.94 | 198.51 | 198.21 | 0.630 |
| 7 | 201.98 | 203.22 | 203.26 | 202.77 | 202.81 | 0.595 |
| 8 | 203.21 | 203.51 | 201.80 | 202.19 | 202.68 | 0.813 |
| 9 | 188.82 | 187.75 | 188.58 | 189.98 | 188.78 | 0.921 |
| 10 | 194.58 | 195.21 | 202.02 | 203.33 | 198.79 | 4.531 |
| 11 | 201.13 | 201.29 | 201.66 | 200.11 | 201.05 | 0.663 |
| 12 | 201.70 | 202.68 | 204.32 | 203.07 | 202.94 | 1.084 |
| 13 | 203.31 | 205.73 | 203.29 | 203.37 | 203.93 | 1.204 |
| 14 | 196.93 | 198.68 | 199.18 | 200.05 | 198.71 | 1.315 |
| 15 | 223.91 | 225.72 | 225.47 | 227.54 | 225.66 | 1.487 |
| 16 | 212.68 | 214.85 | 204.47 | 210.60 | 210.65 | 4.470 |
| 17 | 207.29 | 201.79 | 206.49 | 205.24 | 205.20 | 2.426 |
| 18 | 198.38 | 197.26 | 198.46 | 200.45 | 198.64 | 1.327 |
| 19 | 203.65 | 201.50 | 203.13 | 201.79 | 202.52 | 1.036 |
| 20 | 201.38 | 201.47 | 203.42 | 202.88 | 202.29 | 1.021 |
| 21 | 202.75 | 201.91 | 200.88 | 200.39 | 201.48 | 1.056 |
| 22 | 193.49 | 198.16 | 200.54 | 197.87 | 197.52 | 2.938 |
| 23 | 199.49 | 196.42 | 195.10 | 200.07 | 197.77 | 2.394 |
| 24 | 201.47 | 201.61 | 201.98 | 202.14 | 201.80 | 0.313 |
| 25 | 204.16 | 203.93 | 203.81 | 203.46 | 203.84 | 0.292 |
| 26 | 206.23 | 203.64 | 201.79 | 200.91 | 203.14 | 2.352 |

Lab 4: Invalid data due to noncompliance with Pyraoxystrobin method.

Lab 15: Outlier according to Grubbs Test.

Lab 16: Outlier according to Cochran Test.

Fig.4-1: Pyraoxystrobin SC2 (Batch No: 20130308001)   All labs

Fig.4-2: Pyraoxystrobin SC2 (Batch No: 20130308001)    Labs 4, 15, 16 excluded

Table 5: Pyraoxystrobin SC3 (Batch No: 20130315001)

| Lab | Day 1 |  | Day 2 |  | Mean g/kg | Std. Dev. |
| --- | --- | --- | --- | --- | --- | --- |
|  | A | B | A | B |  |  |
| 1 | 213.82 | 216.00 | 214.23 | 214.06 | 214.53 | 0.996 |
| 2 | 200.98 | 208.32 | 202.00 | 214.82 | 206.53 | 6.410 |
| 3 | 210.28 | 209.84 | 210.18 | 210.61 | 210.23 | 0.317 |
| 4 | 228.86 | 210.76 | 204.61 | 225.94 | 217.54 | 11.717 |
| 5 | 205.09 | 206.83 | 206.30 | 205.98 | 206.05 | 0.730 |
| 6 | 203.83 | 203.59 | 204.13 | 203.93 | 203.87 | 0.224 |
| 7 | 208.49 | 207.50 | 209.28 | 209.38 | 208.66 | 0.871 |
| 8 | 208.07 | 208.78 | 207.53 | 206.90 | 207.82 | 0.799 |
| 9 | 190.86 | 190.39 | 192.02 | 191.94 | 191.30 | 0.806 |
| 10 | 202.86 | 202.00 | 208.51 | 207.93 | 205.33 | 3.370 |
| 11 | 205.71 | 205.88 | 204.73 | 204.30 | 205.16 | 0.763 |
| 12 | 204.93 | 205.21 | 207.44 | 206.46 | 206.01 | 1.162 |
| 13 | 208.88 | 210.86 | 210.22 | 208.49 | 209.61 | 1.114 |
| 14 | 203.14 | 203.08 | 203.00 | 204.84 | 203.52 | 0.885 |
| 15 | 225.69 | 222.35 | 226.18 | 222.89 | 224.28 | 1.937 |
| 16 | 213.69 | 215.12 | 214.34 | 208.94 | 213.02 | 2.784 |
| 17 | 210.69 | 208.17 | 209.90 | 210.39 | 209.79 | 1.126 |
| 18 | 202.39 | 202.59 | 204.14 | 204.86 | 203.50 | 1.200 |
| 19 | 208.79 | 205.18 | 207.49 | 207.61 | 207.27 | 1.510 |
| 20 | 206.36 | 205.75 | 207.85 | 208.81 | 207.19 | 1.393 |
| 21 | 207.55 | 208.00 | 205.07 | 204.42 | 206.26 | 1.779 |
| 22 | 204.24 | 203.82 | 210.48 | 207.36 | 206.48 | 3.102 |
| 23 | 209.30 | 204.71 | 181.04 | 195.76 | 197.70 | 12.450 |
| 24 | 207.19 | 207.26 | 207.51 | 207.68 | 207.41 | 0.226 |
| 25 | 201.45 | 201.57 | 200.17 | 200.35 | 200.89 | 0.727 |
| 26 | 209.88 | 207.16 | 209.36 | 207.54 | 208.49 | 1.337 |

Lab 4: Invalid data due to noncompliance with Pyraoxystrobin method.

Lab 2 and Lab 23: Outlier according to Cochran Test.

Lab 9 and 15: Straggler according to Grubbs Test.

Fig.5-1: Pyraoxystrobin SC3 (Batch No: 20130315001)   All labs

Fig.5-2: Pyraoxystrobin SC3 (Batch No: 20130315001)    Labs 2, 4, 9, 15 and 23 excluded

# Summary of the results

Table 6   Summary of the results of all laboratories

|  | TC-1 | TC-2 | SC-1 | SC-2 | SC-3 |
| --- | --- | --- | --- | --- | --- |
| x | 973.45 | 973.95 | 202.81 | 202.93 | 206.83 |
| L | 25 | 25 | 25 | 25 | 25 |
| Sr | 7.654 | 8.439 | 2.182 | 2.028 | 3.154 |
| SR | 12.006 | 13.193 | 6.381 | 6.680 | 6.501 |
| r | 21.430 | 23.629 | 6.110 | 5.679 | 8.831 |
| R | 33.617 | 36.942 | 17.866 | 18.704 | 18.203 |
| RSDr | 0.786 | 0.866 | 1.076 | 0.999 | 1.525 |
| RSDR | 1.233 | 1.355 | 3.146 | 3.292 | 3.143 |
| RSDR(Hor) | 2.008 | 2.008 | 2.543 | 2.543 | 2.535 |

( values given in units of g/kg )

Table 7   Summary of the results after elimination of outlier values

|  | TC-1 | TC-2 | SC-1 | SC-2 | SC-3 |
| --- | --- | --- | --- | --- | --- |
| x | 972.36 | 971.62 | 202.86 | 201.61 | 207.19 |
| L | 22 | 22 | 22 | 23 | 21 |
| Sr | 5.363 | 4.651 | 1.740 | 1.873 | 1.515 |
| SR | 7.849 | 8.184 | 4.653 | 4.464 | 3.450 |
| r | 15.018 | 13.024 | 4.872 | 5.243 | 4.243 |
| R | 21.978 | 22.916 | 13.029 | 12.499 | 9.660 |
| RSDr | 0.552 | 0.479 | 0.858 | 0.929 | 0.731 |
| RSDR | 0.807 | 0.842 | 2.294 | 2.214 | 1.665 |
| RSDR(Hor) | 2.008 | 2.009 | 2.543 | 2.545 | 2.535 |

( values given in units of g/kg )

Where:

x             = average, in unit of g/kg

L             = number of laboratories

Sr            = repeatability standard deviation

SR            = reproducibility standard deviation =(Sr2+SL2)

r             = repeatability ( Sr·2.8 )

R             = reproducibility ( SR·2.8 )

RSDr         = repeatability relative standard deviation ( 100·Sr/x )

RSDR         = reproducibility relative standard deviation ( 100·SR/x )

RSDR(Hor)     = Horwitz value calculated from: 2(1-0.5log c)

where c        = the concentration of the analyte as a decimal fraction

# Statistical formulas

Yi    = mean of the various laboratories

Si    = standard deviation

P    = number of laboratories

n    = number of measurements ( here n=4 )

T1=i=1pYi

T2=i=1pYi2

T3=i=1pSi2

Repeatability and reproducibility were calculated as follows:

Sr2=T3P

SL2=PT2−T12P(P-1)−Sr2n

SR2=Sr2+SL2

r=2.8∗Sr2

R=2.8∗SR2

# Discussion

Following the successful outcome of the full scale collaborative study organized by SYRICI, an international CIPAC collaborative study was initiated in November 2013 to test a specific HPLC method for the determination of pyraoxystrobin.

26 laboratories had announced to participate the CIPAC trial and sent back results.

The data from each of the laboratories were reviewed to determine if there were any problems with analysis procedure used, chromatography or reporting results, which might affect the analyses results. The changes, deviations, and observations which were noted will not be expected to affect the analyses results significantly.

But there were some problems in Lab 4. The response factor for the calibration solution B differed much more than 1.0% from that for the calibration solution A, which did not comply with the procedure of HPLC method of pyraoxystrobin for CIPAC Full Scale Collaborative Trial. So the data of the Lab 4 was invalid and excluded from the statistical evaluation.

If the results of 25 laboratories (excluding the invalid data of Lab 4) participated in the collaborative trial are taken into account for the statistical evaluation, i.e. all stragglers and outliers according to Cochran test and Grubbs test are left in the evaluation and no data are rejected, the Horwitz criterion will be fulfilled in case of TC1, TC2 and not fulfilled in case of SC1, SC2, SC3. (table 6)

The Horwitz criterion is improved for TC1 after elimination of two outliers according to Cochran test (Lab 13 and Lab 17) and one outlier according to Grubbs test (Lab 16).

The Horwitz criterion is improved for TC2 after elimination of two outliers according to Cochran test (Lab 12 and Lab 17) and one outlier according to Grubbs test (Lab 16).

The Horwitz criterion is fulfilled for SC1 after elimination of one outlier according to Cochran test (Lab 23), one outlier according to Dixon test (Lab 9) and one outlier due to deviation from the method (Lab 16).

The Horwitz criterion is fulfilled for SC2 after elimination of one outlier according to Cochran test (Lab 16) and one outlier according to Grubbs test (Lab 15).

The Horwitz criterion is fulfilled for SC3 after elimination of two outliers according to Cochran test (Lab 2 and Lab 23) and two stragglers according to Grubbs test (Lab 9 and Lab 15).

Overview: outliers and stragglers identified and allocated to the participant

| Sample No. | Lab ID No. Identification of outliers and stragglers |
| --- | --- |
| TC1 | 13, 16, 17 |
| TC2 | 12, 16, 17 |
| SC1 | 9, 16, 23 |
| SC2 | 15, 16 |
| SC3 | 2, 9, 15, 23 |

Conclusion:

We would like to propose the analytical method for Pyraoxystrobin to become provisional.

Acknowledgements:

The organizer wishes to thank all laboratories and their staff who participated in this study.