## Slide 1: 28-Homobrassinolide- CIPAC Small Scale Collaborative Study
CIPAC technical meeting 2020
Jason Zhang (Chinese Pesticide Advisory Committee, CHIPAC)
Method developed by Jiangxi Windeal Biotechnology Co., Ltd.

## Slide 2: GENERAL INFORMATION
Common name | 28-Homobrassinolide
Chemical Name | (5S,6R)-10-((2S,3R,4R,5S)-5-ethyl-3,4-dihydroxy-6-methylheptan-2-yl)-5,6-dihydroxy-7a,9a-dimethyltetradecahydro-1H-benzo[c]indeno[5,4-e]oxepin-3(12bH)-one
Empirical formula | C29H50O6
Structure | 
RMM | 494.8
m.p. | 256~257°C
Solubility (g/L, 20℃) | In water 5mg/l, acetonitrile 1.3g/l, ethanol 5.2g/l, methanol 2.7g/l at 20 °C.
Description | White powder
Stability | Stable in neutral and weak alkaline condition but hydrolysed in acidic conditions
Formulation | Emulsifiable concentrate, soluble liquid

## Slide 3: PARTICIPANTS
Two technical samples, two SL samples and one EC sample were sent to the following 6 participants in February 2020.
By the end of April 2020, all 6 participants provided their results to CHIPAC.
Lab No. | Name of responsible person | Lab Name | Country
1 | Liu ShuZhen | Jiangxi Windeal Biotechnology Co., Ltd. - Quality Department | Jiangxi, China
2 | Sun Fengying | Kaifeng Yitian Biotechnology Co., Ltd. - quality control department | Henan, China
3 | Chen Mirror | GreenTech Laboratory Co., Ltd. | Shanghai, China
4 | Mei Quanfu | Jiangxi Buffett Chemical Co., Ltd. - quality control department | Jiangxi, China
5 | Chang Feng | Shandong Huihan Biotechnology Co., Ltd. - analysis room | Shandong, China
6 | Liu Xinsheng | Guangdong Zhongxun Agricultural Technology Co., Ltd. - PTD Center | Guangdong, China

## Slide 4
 | TC-1 | TC-2 | SL-1 | SL-2 | EC
Xm | 955.7 | 950.7 | 0.041 | 0.042 | 0.100
L | 6 | 6 | 6 | 6 | 6
Sr | 1.862 | 2.535 | 0.00053 | 0.00098 | 0.00101
SR | 2.422 | 6.516 | 0.00146 | 0.00229 | 0.00715
r | 5.214 | 7.098 | 0.00148 | 0.00274 | 0.00283
R | 6.782 | 18.245 | 0.00409 | 0.00641 | 0.02002
RSDr | 0.195 | 0.267 | 1.296 | 2.342 | 1.017
RSDR | 0.253 | 0.685 | 3.585 | 5.485 | 7.173
RSDR (Hor) | 2.014 | 2.015 | 9.160 | 9.123 | 8.004
HorRat Value | 0.126 | 0.340 | 0.391 | 0.601 | 0.896
Summary of the statistical evaluation no elimination of any outliers
Xm = average
L = number of laboratories
Sr = repeatability standard deviation
SR = reproducibility standard deviation
RSDr = repeatability relative standard deviation
RSDR = reproducibility relative standard deviation
r = repeatability
R = reproducibility
RSDR (Hor) = Horwitz value calculated from: 2^(1 - 0.5log c) where c = the concentration of the analyte as a decimal fraction
28-Homobrassinolide- CIPAC SMALL SCALE COLLABORATIVE STUDY

## Slide 5
 | TC-1 | TC-2 | SL-1 | SL-2 | EC
Xm | 955.7 | 953.1 | 0.041 | 0.042 | 0.102
L | 6 | 5 | 6 | 6 | 5
Sr | 1.862 | 2.750 | 0.00053 | 0.00098 | 0.00109
SR | 2.422 | 2.786 | 0.00146 | 0.00229 | 0.00121
r | 5.214 | 7.7 | 0.00148 | 0.00274 | 0.00305
R | 6.782 | 7.801 | 0.00409 | 0.00641 | 0.00339
RSDr | 0.195 | 0.288 | 1.296 | 2.342 | 1.061
RSDR | 0.253 | 0.292 | 3.585 | 5.485 | 1.179
RSDR (Hor) | 2.014 | 2.014 | 9.160 | 9.123 | 7.970
HorRat Value | 0.126 | 0.145 | 0.391 | 0.601 | 0.148
Summary of the statistical evaluation with elimination of stragglers
Xm = average
L = number of laboratories
Sr = repeatability standard deviation
SR = reproducibility standard deviation
RSDr = repeatability relative standard deviation
RSDR = reproducibility relative standard deviation
r = repeatability
R = reproducibility
RSDR (Hor) = Horwitz value calculated from: 2^(1 - 0.5log c) where c = the concentration of the analyte as a decimal fraction

## Slide 6
 | TC-1 | TC-2 | SL-1 | SL-2 | EC
Xm | 955.7 | 953.1 | 0.041 | 0.042 | 0.102
L | 6 | 5 | 5 | 6 | 5
Sr | 1.862 | 2.750 | 0.00058 | 0.00098 | 0.00109
SR | 2.422 | 2.786 | 0.00072 | 0.00229 | 0.00121
r | 5.214 | 7.7 | 0.00162 | 0.00274 | 0.00305
R | 6.782 | 7.801 | 0.00202 | 0.00641 | 0.00339
RSDr | 0.195 | 0.288 | 1.401 | 2.342 | 1.061
RSDR | 0.253 | 0.292 | 1.738 | 5.485 | 1.179
RSDR (Hor) | 2.014 | 2.014 | 9.142 | 9.123 | 7.970
HorRat Value | 0.126 | 0.145 | 0.190 | 0.601 | 0.148
Summary of the statistical evaluation with elimination of outliers
Xm = average
L = number of laboratories
Sr = repeatability standard deviation
SR = reproducibility standard deviation
RSDr = repeatability relative standard deviation
RSDR = reproducibility relative standard deviation
r = repeatability
R = reproducibility
RSDR (Hor) = Horwitz value calculated from: 2^(1 - 0.5log c) where c = the concentration of the analyte as a decimal fraction

## Slide 7
Analytical Method
Outline of Method
The 28-Homobrassinolide content of the samples is determined by high performance liquid chromatography on ODS-C18 film stainless column with UV detector at 220 nm, quantified by external standard method.
Column: | 250mm x 4.6 mm (id), packed with ODS-C18, or equivalent | Eluent: | acetonitrile+ water = 80 + 20（v/v）
Column Temperature: | 25℃ | Flow Rate: | 1.0 ml/min
Detector Wavelength： | 220nm | Injection Volume: | 10μL
Retention time: | Approximately 20.5min |  | 

## Slide 8
Analytical Method
Preparation of the test substance solution and determination
Preparation of standard solution: prepare standard solution in duplicate. Weigh 0.015g (to the nearest 0.1mg) 28-Homobrassinolide standard into 25ml volumetric flask, dissolved by 10ml methanol. Add 10ml phenylboronic acid solution, react 30min in calorstat at 50°C. Allow the solution to cool to ambient temperature and fill to the mark with methanol. Mix thoroughly and place the flask in an ultrasonic bath for 5 min, then filter the solution through a 0.45μm filter membrane prior to analysis.
Preparation of sample solution: prepare sample solution in duplicate. Weigh 0.015g (to the nearest 0.1mg) sufficient sample to contain about 15mg 28-Homobrassinolide into 25ml volumetric flask, dissolved by 10ml methanol. Add 10ml phenylboronic acid solution, react 30min in calorstat at 50°C. Allow the solution to cool to ambient temperature and fill to the mark with methanol. Mix thoroughly and place the flask in an ultrasonic bath for 5 min, then filter the solution through a 0.45μm filter membrane prior to analysis. (Sample solutions S1 and S2).
Determination: Inject in duplicate 10μL portions of each sample solution bracketing them by injections of the calibration solutions as follows: CA, S1, S1, CB, S2, S2, CA, etc.

## Slide 9
Study format
Samples
28-Homobrassinolide Technical (SAMPLE A)
28-Homobrassinolide Technical (SAMPLE B)
28-Homobrassinolide Soluble Liquid (SAMPLE C)
28-Homobrassinolide Soluble Liquid (SAMPLE D)
28-Homobrassinolide Emulsifiable Concentrate (SAMPLE E)
Protocol
The samples were analyzed on two different days with duplicate injections weighting per sample. Test and calibration solutions were prepared fresh on each day. The sample content was calculated using the mean value of the duplicate injections.

## Slide 10
Analytical conditions
Lab Number | Analytical Conditions
Laboratory 1 | Column: SHIMADZU, 4.6mm×250mm, 8FR98171 Remarks: None
Laboratory 2 | Column: Elite, 4.6mm×250mm, E2616571 Remarks: None
Laboratory 3 | Column: Shimadzu InerStainTM C18, 250mm×4.6mm, 5μm Remarks: None
Laboratory 4 | Column: Agilent, 4.6mm×250mm, USNH075205 Remarks: None
Laboratory 5 | Column: Agilent, TC-C18(2), 250mm×4.6mm Remarks: Column temperature is 35℃, Injection volume is 5µL.
Laboratory 6 | Column: XTERRA MS C18, 4.6mm×250mm, 03273926014075 Remarks: None

## Slide 11
DATA Evaluation and Discussion
ORINGINAL DATA (g/kg)
 | 28-Homobrassinolide SAMPLE A |  |  |  | 28-Homobrassinolide SAMPLE B |  |  |  | 28-Homobrassinolide SAMPLE C | 
 | Day1 |  | Day2 |  | Day1 |  | Day2 |  | Day1 | 
Laboratory 1 | 959.5 | 957.7 | 955.1 | 956.1 | 951.6 | 952.3 | 951.0 | 952.9 | 0.041 | 0.041
Laboratory 2 | 959.7 | 957.1 | 958.7 | 954.6 | 950.5 | 950.8 | 953.4 | 959.2 | 0.040 | 0.042
Laboratory 3 | 953.1 | 956.4 | 953.5 | 954.8 | 937.8 | 939.7 | 938.4 | 937.9 | 0.038 | 0.038
Laboratory 4 | 954.6 | 960.5 | 954.2 | 957.9 | 954.6 | 951.6 | 952.8 | 952.1 | 0.042 | 0.042
Laboratory 5 | 952.0 | 952.0 | 953.2 | 954.1 | 957.9 | 956.5 | 952.1 | 955.4 | 0.040 | 0.041
Laboratory 6 | 955.6 | 956.1 | 955.7 | 955.7 | 953.2 | 948.8 | 949.5 | 956.6 | 0.042 | 0.042

## Slide 12
 | 28-Homobrassinolide SAMPLE C |  | 28-Homobrassinolide SAMPLE D |  |  |  | 28-Homobrassinolide SAMPLE E |  |  | 
 | Day2 |  | Day1 |  | Day2 |  | Day1 |  | Day2 | 
Laboratory 1 | 0.041 | 0.041 | 0.042 | 0.042 | 0.042 | 0.042 | 0.101 | 0.101 | 0.103 | 0.101
Laboratory 2 | 0.041 | 0.040 | 0.046 | 0.047 | 0.044 | 0.042 | 0.104 | 0.105 | 0.101 | 0.102
Laboratory 3 | 0.038 | 0.038 | 0.038 | 0.039 | 0.038 | 0.038 | 0.085 | 0.085 | 0.086 | 0.0854
Laboratory 4 | 0.041 | 0.042 | 0.043 | 0.042 | 0.042 | 0.043 | 0.103 | 0.101 | 0.103 | 0.102
Laboratory 5 | 0.041 | 0.041 | 0.041 | 0.041 | 0.041 | 0.041 | 0.103 | 0.104 | 0.104 | 0.103
Laboratory 6 | 0.042 | 0.041 | 0.042 | 0.042 | 0.043 | 0.042 | 0.103 | 0.102 | 0.103 | 0.102

## Slide 13: DATA EVALUATION AND DISCUSSION
MEAN VALUES (g/kg)
 | 28-Homobrassinolide SAMPLE A | 28-Homobrassinolide SAMPLE B | 28-Homobrassinolide SAMPLE C | 28-Homobrassinolide SAMPLE D | 28-Homobrassinolide SAMPLE E
Laboratory 1 | 957.1 | 952.0 | 0.041 | 0.042 | 0.102
Laboratory 2 | 957.5 | 953.5 | 0.041 | 0.045 | 0.103
Laboratory 3 | 954.5 | 938.5+/++ | 0.038+ | 0.038 | 0.086+/++
Laboratory 4 | 956.9 | 952.8 | 0.042 | 0.042 | 0.102
Laboratory 5 | 952.9 | 955.5 | 0.041 | 0.041 | 0.104
Laboratory 6 | 955.8 | 952.0 | 0.041 | 0.042 | 0.102
+ Gubbs Test straggler
++ Gubbs Test outlier

## Slide 14
DATA Evaluation and Discussion
Xm | 955.7
L | 6
Sr | 1.862
SR | 2.422
r | 5.214
R | 6.782
RSDr | 0.195
RSDR | 0.253
RSDR(Hor) | 2.014
HorRat | 0.126
Fig. 1 – 28-Homobrassinolide Sample A with outliers /stragglers

## Slide 15
Xm | 950.7
L | 6
Sr | 2.535
SR | 6.516
r | 7.098
R | 18.245
RSDr | 0.267
RSDR | 0.685
RSDR(Hor) | 2.015
HorRat | 0.340
Fig. 2 – 28-Homobrassinolide Sample B with outliers /stragglers

## Slide 16
Xm | 0.041
L | 6
Sr | 0.00053
SR | 0.00146
r | 0.00148
R | 0.00409
RSDr | 1.296
RSDR | 3.585
RSDR(Hor) | 9.160
HorRat | 0.391
Fig. 3 – 28-Homobrassinolide Sample C with outliers /stragglers

## Slide 17
Xm | 0.042
L | 6
Sr | 0.00098
SR | 0.00229
r | 0.00274
R | 0.00641
RSDr | 2.342
RSDR | 5.485
RSDR(Hor) | 9.123
HorRat | 0.601
Fig. 4 – 28-Homobrassinolide Sample D with outliers /stragglers

## Slide 18
Xm | 0.100
L | 6
Sr | 0.00101
SR | 0.00715
r | 0.00283
R | 0.02002
RSDr | 1.017
RSDR | 7.173
RSDR(Hor) | 8.004
HorRat | 0.896
Fig. 5 – 28-Homobrassinolide Sample E with outliers /stragglers

## Slide 19
Summary
Summary of the statistical evaluation with outliers /stragglers
Xm = average
L = number of laboratories
Sr = repeatability standard deviation
SR = reproducibility standard deviation
RSDr = repeatability relative standard deviation
RSDR = reproducibility relative standard deviation
r = repeatability
R = reproducibility
RSDR (Hor) = Horwitz value calculated from: 2^(1 - 0.5log c) where c = the concentration of the analyte as a decimal fraction
 | sample A | sample B | sample C | sample D | sample E
Xm | 955.7 | 950.7 | 0.041 | 0.042 | 0.100
L | 6 | 6 | 6 | 6 | 6
Sr | 1.862 | 2.535 | 0.00053 | 0.00098 | 0.00101
SR | 2.422 | 6.516 | 0.00146 | 0.00229 | 0.00715
r | 5.214 | 7.098 | 0.00148 | 0.00274 | 0.00283
R | 6.782 | 18.245 | 0.00409 | 0.00641 | 0.02002
RSDr | 0.195 | 0.267 | 1.296 | 2.342 | 1.017
RSDR | 0.253 | 0.685 | 3.585 | 5.485 | 7.173
RSDR(Hor) | 2.014 | 2.015 | 9.160 | 9.123 | 8.004
HorRat | 0.126 | 0.340 | 0.391 | 0.601 | 0.896

## Slide 20
 | sample A | sample B | sample C | sample D | sample E
Xm | 955.7 | 953.1 | 0.041 | 0.042 | 0.102
L | 6 | 5 | 6 | 6 | 5
Sr | 1.862 | 2.750 | 0.00053 | 0.00098 | 0.00109
SR | 2.422 | 2.786 | 0.00146 | 0.00229 | 0.00121
r | 5.214 | 7.7 | 0.00148 | 0.00274 | 0.00305
R | 6.782 | 7.801 | 0.00409 | 0.00641 | 0.00339
RSDr | 0.195 | 0.288 | 1.296 | 2.342 | 1.061
RSDR | 0.253 | 0.292 | 3.585 | 5.485 | 1.179
RSDR(Hor) | 2.014 | 2.014 | 9.160 | 9.123 | 7.970
HorRat | 0.126 | 0.145 | 0.391 | 0.601 | 0.148
Sample B Results of Lab 3 eliminated, Sample E Results of Lab 3 eliminated.
Xm = average
L = number of laboratories
Sr = repeatability standard deviation
SR = reproducibility standard deviation
RSDr = repeatability relative standard deviation
RSDR = reproducibility relative standard deviation
r = repeatability
R = reproducibility
RSDR (Hor) = Horwitz value calculated from: 2^(1 - 0.5log c) where c = the concentration of the analyte as a decimal fraction
Summary of the statistical evaluation without stragglers

## Slide 21
 | sample A | sample B | sample C | sample D | sample E
Xm | 955.7 | 953.1 | 0.041 | 0.042 | 0.102
L | 6 | 5 | 5 | 6 | 5
Sr | 1.862 | 2.750 | 0.00058 | 0.00098 | 0.00109
SR | 2.422 | 2.786 | 0.00072 | 0.00229 | 0.00121
r | 5.214 | 7.7 | 0.00162 | 0.00274 | 0.00305
R | 6.782 | 7.801 | 0.00202 | 0.00641 | 0.00339
RSDr | 0.195 | 0.288 | 1.401 | 2.342 | 1.061
RSDR | 0.253 | 0.292 | 1.738 | 5.485 | 1.179
RSDR(Hor) | 2.014 | 2.014 | 9.142 | 9.123 | 7.970
HorRat | 0.126 | 0.145 | 0.190 | 0.601 | 0.148
Sample B Results of Lab 3 eliminated, Sample C Results of Lab 3 eliminated, Sample E Results of Lab 3 eliminated.
Xm = average
L = number of laboratories
Sr = repeatability standard deviation
SR = reproducibility standard deviation
RSDr = repeatability relative standard deviation
RSDR = reproducibility relative standard deviation
r = repeatability
R = reproducibility
RSDR (Hor) = Horwitz value calculated from: 2^(1 - 0.5log c) where c = the concentration of the analyte as a decimal fraction
Summary of the statistical evaluation without outliers

## Slide 22
conclusions
CHIPAC proposes to proceed with a large scale collaborative study.
For all samples, the values of RSDR (reproducibility relative standard deviation) were less than Horwitz’s value. As a reference, all HorRat values were not greater than 1.0. The proposed method is considered to be appropriate for the determination of 28-Homobrassinolide in technical material, SL and EC formulation.

## Slide 23
Thanks for your attention!